## papers_passed
Reference | Index | PMID | Title | Authors | Citation | First Author | Journal/Book | Publication Year | Create Date | PMCID | NIHMS ID | DOI
Hossain22 | 8 | 35590859 | Motion Artifacts Correction from Single-Channel EEG and fNIRS Signals Using Novel Wavelet Packet Decomposition in Combination with Canonical Correlation Analysis | Hossain MS, Chowdhury MEH, Reaz MBI, Ali SHM, Bakar AAA, Kiranyaz S, Khandakar A, Alhatou M, Habib R, Hossain MM. | Sensors (Basel). 2022 Apr 21;22(9):3169. doi: 10.3390/s22093169. | Hossain MS | Sensors (Basel) | 2022 | 2022-05-20 | PMC9102309 |  | 10.3390/s22093169
Abu22 | 10 | 35459033 | Improved Cognitive Vigilance Assessment after Artifact Reduction with Wavelet Independent Component Analysis | Abu Farha N, Al-Shargie F, Tariq U, Al-Nashash H. | Sensors (Basel). 2022 Apr 15;22(8):3051. doi: 10.3390/s22083051. | Abu Farha N | Sensors (Basel) | 2022 | 2022-04-23 | PMC9033092 |  | 10.3390/s22083051
Flo22 | 16 | 35093730 | Automated Pipeline for Infants Continuous EEG (APICE): A flexible pipeline for developmental cognitive studies | Fló A, Gennari G, Benjamin L, Dehaene-Lambertz G. | Dev Cogn Neurosci. 2022 Apr;54:101077. doi: 10.1016/j.dcn.2022.101077. Epub 2022 Jan 25. | Fló A | Dev Cogn Neurosci | 2022 | 2022-01-30 | PMC8804179 |  | 10.1016/j.dcn.2022.101077
Kumaravel22 | 17 | 35085870 | NEAR: An artifact removal pipeline for human newborn EEG data | Kumaravel VP, Farella E, Parise E, Buiatti M. | Dev Cogn Neurosci. 2022 Apr;54:101068. doi: 10.1016/j.dcn.2022.101068. Epub 2022 Jan 15. | Kumaravel VP | Dev Cogn Neurosci | 2022 | 2022-01-27 | PMC8800139 |  | 10.1016/j.dcn.2022.101068
Soghoyan21 | 35 | 34924988 | A Toolbox and Crowdsourcing Platform for Automatic Labeling of Independent Components in Electroencephalography | Soghoyan G, Ledovsky A, Nekrashevich M, Martynova O, Polikanova I, Portnova G, Rebreikina A, Sysoeva O, Sharaev M. | Front Neuroinform. 2021 Dec 2;15:720229. doi: 10.3389/fninf.2021.720229. eCollection 2021. | Soghoyan G | Front Neuroinform | 2021 | 2021-12-20 | PMC8675888 |  | 10.3389/fninf.2021.720229
Karpiel21 | 40 | 34833790 | The Influence of Filters on EEG-ERP Testing: Analysis of Motor Cortex in Healthy Subjects | Karpiel I, Kurasz Z, Kurasz R, Duch K. | Sensors (Basel). 2021 Nov 19;21(22):7711. doi: 10.3390/s21227711. | Karpiel I | Sensors (Basel) | 2021 | 2021-11-27 | PMC8619013 |  | 10.3390/s21227711
Beach21 | 47 | 34584747 | Motion artefact removal in electroencephalography and electrocardiography by using multichannel inertial measurement units and adaptive filtering | Beach C, Li M, Balaban E, Casson AJ. | Healthc Technol Lett. 2021 Jun 24;8(5):128-138. doi: 10.1049/htl2.12016. eCollection 2021 Oct. | Beach C | Healthc Technol Lett | 2021 | 2021-09-29 | PMC8450177 |  | 10.1049/htl2.12016
Dai21 | 52 | 34573795 | A Fast Approach to Removing Muscle Artifacts for EEG with Signal Serialization Based Ensemble Empirical Mode Decomposition | Dai Y, Duan F, Feng F, Sun Z, Zhang Y, Caiafa CF, Marti-Puig P, Solé-Casals J. | Entropy (Basel). 2021 Sep 6;23(9):1170. doi: 10.3390/e23091170. | Dai Y | Entropy (Basel) | 2021 | 2021-09-28 | PMC8465074 |  | 10.3390/e23091170
Tost21 | 54 | 34441170 | Choosing Strategies to Deal with Artifactual EEG Data in Children with Cognitive Impairment | Tost A, Migliorelli C, Bachiller A, Medina-Rivera I, Romero S, García-Cazorla Á, Mañanas MA. | Entropy (Basel). 2021 Aug 11;23(8):1030. doi: 10.3390/e23081030. | Tost A | Entropy (Basel) | 2021 | 2021-08-27 | PMC8392530 |  | 10.3390/e23081030
Kaur21 | 56 | 35154587 | EEG Artifact Removal System for Depression Using a Hybrid Denoising Approach | Kaur C, Singh P, Sahni S. | Basic Clin Neurosci. 2021 Jul-Aug;12(4):465-476. doi: 10.32598/bcn.2021.1388.2. Epub 2021 Jul 1. | Kaur C | Basic Clin Neurosci | 2021 | 2022-02-14 | PMC8817173 |  | 10.32598/bcn.2021.1388.2
Rosanne21 | 73 | 33897342 | Adaptive Filtering for Improved EEG-Based Mental Workload Assessment of Ambulant Users | Rosanne O, Albuquerque I, Cassani R, Gagnon JF, Tremblay S, Falk TH. | Front Neurosci. 2021 Apr 7;15:611962. doi: 10.3389/fnins.2021.611962. eCollection 2021. | Rosanne O | Front Neurosci | 2021 | 2021-04-26 | PMC8058356 |  | 10.3389/fnins.2021.611962
Nahmias21 | 82 | 33642972 | Quantifying Signal Quality From Unimodal and Multimodal Sources: Application to EEG With Ocular and Motion Artifacts | Nahmias DO, Kontson KL. | Front Neurosci. 2021 Feb 12;15:566004. doi: 10.3389/fnins.2021.566004. eCollection 2021. | Nahmias DO | Front Neurosci | 2021 | 2021-03-01 | PMC7906969 |  | 10.3389/fnins.2021.566004
Saba-Sadiya21 | 86 | 34713069 | Unsupervised EEG Artifact Detection and Correction | Saba-Sadiya S, Chantland E, Alhanai T, Liu T, Ghassemi MM. | Front Digit Health. 2021 Jan 22;2:608920. doi: 10.3389/fdgth.2020.608920. eCollection 2020. | Saba-Sadiya S | Front Digit Health | 2021 | 2021-10-29 | PMC8521924 |  | 10.3389/fdgth.2020.608920
Li21 | 89 | 33584176 | Electromyogram (EMG) Removal by Adding Sources of EMG (ERASE)-A Novel ICA-Based Algorithm for Removing Myoelectric Artifacts From EEG | Li Y, Wang PT, Vaidya MP, Flint RD, Liu CY, Slutzky MW, Do AH. | Front Neurosci. 2021 Jan 15;14:597941. doi: 10.3389/fnins.2020.597941. eCollection 2020. | Li Y | Front Neurosci | 2021 | 2021-02-15 | PMC7873899 |  | 10.3389/fnins.2020.597941
Anders20 | 106 | 32860085 | The influence of motor tasks and cut-off parameter selection on artifact subspace reconstruction in EEG recordings | Anders P, Müller H, Skjæret-Maroni N, Vereijken B, Baumeister J. | Med Biol Eng Comput. 2020 Nov;58(11):2673-2683. doi: 10.1007/s11517-020-02252-3. Epub 2020 Aug 28. | Anders P | Med Biol Eng Comput | 2020 | 2020-08-30 | PMC7560919 |  | 10.1007/s11517-020-02252-3
Zhang20 | 117 | 32943621 | Fully organic compliant dry electrodes self-adhesive to skin for long-term motion-robust epidermal biopotential monitoring | Zhang L, Kumar KS, He H, Cai CJ, He X, Gao H, Yue S, Li C, Seet RC, Ren H, Ouyang J. | Nat Commun. 2020 Sep 17;11(1):4683. doi: 10.1038/s41467-020-18503-8. | Zhang L | Nat Commun | 2020 | 2020-09-18 | PMC7499260 |  | 10.1038/s41467-020-18503-8
Leach20 | 126 | 32185818 | Adjusting ADJUST: Optimizing the ADJUST algorithm for pediatric data using geodesic nets | Leach SC, Morales S, Bowers ME, Buzzell GA, Debnath R, Beall D, Fox NA. | Psychophysiology. 2020 Aug;57(8):e13566. doi: 10.1111/psyp.13566. Epub 2020 Mar 17. | Leach SC | Psychophysiology | 2020 | 2020-03-19 | PMC7402217 | NIHMS1611442 | 10.1111/psyp.13566
Saini20 | 139 | 32431850 | Effective automated method for detection and suppression of muscle artefacts from single-channel EEG signal | Saini M, Satija U, Upadhayay MD. | Healthc Technol Lett. 2020 Apr 14;7(2):35-40. doi: 10.1049/htl.2019.0053. eCollection 2020 Apr. | Saini M | Healthc Technol Lett | 2020 | 2020-05-21 | PMC7199290 |  | 10.1049/htl.2019.0053
Mucarquer20 | 144 | 32205896 | Improving EEG Muscle Artifact Removal With an EMG Array | Mucarquer JA, Prado P, Escobar MJ, El-Deredy W, Zañartu M. | IEEE Trans Instrum Meas. 2020 Mar;69(3):815-824. doi: 10.1109/tim.2019.2906967. Epub 2019 May 1. | Mucarquer JA | IEEE Trans Instrum Meas | 2020 | 2020-03-25 | PMC7088455 | NIHMS1069549 | 10.1109/tim.2019.2906967
Nordin20 | 145 | 31199248 | Faster Gait Speeds Reduce Alpha and Beta EEG Spectral Power From Human Sensorimotor Cortex | Nordin AD, Hairston WD, Ferris DP. | IEEE Trans Biomed Eng. 2020 Mar;67(3):842-853. doi: 10.1109/TBME.2019.2921766. Epub 2019 Jun 13. | Nordin AD | IEEE Trans Biomed Eng | 2020 | 2019-06-15 | PMC7134343 | NIHMS1565487 | 10.1109/TBME.2019.2921766
Liu19 | 156 | 31976053 | Removal of EMG Artifacts from Multichannel EEG Signals Using Combined Singular Spectrum Analysis and Canonical Correlation Analysis | Liu Q, Liu A, Zhang X, Chen X, Qian R, Chen X. | J Healthc Eng. 2019 Dec 30;2019:4159676. doi: 10.1155/2019/4159676. eCollection 2019. | Liu Q | J Healthc Eng | 2019 | 2020-01-25 | PMC6955116 |  | 10.1155/2019/4159676
Benda19 | 162 | 31795398 | Peak Detection with Online Electroencephalography (EEG) Artifact Removal for Brain-Computer Interface (BCI) Purposes | Benda M, Volosyak I. | Brain Sci. 2019 Nov 29;9(12):347. doi: 10.3390/brainsci9120347. | Benda M | Brain Sci | 2019 | 2019-12-05 | PMC6955994 |  | 10.3390/brainsci9120347
Val-Calvo19 | 163 | 31849630 | Optimization of Real-Time EEG Artifact Removal and Emotion Estimation for Human-Robot Interaction Applications | Val-Calvo M, Álvarez-Sánchez JR, Ferrández-Vicente JM, Fernández E. | Front Comput Neurosci. 2019 Nov 26;13:80. doi: 10.3389/fncom.2019.00080. eCollection 2019. | Val-Calvo M | Front Comput Neurosci | 2019 | 2019-12-19 | PMC6889828 |  | 10.3389/fncom.2019.00080
Saavedra19 | 171 | 31534448 | Wavelet-Based Semblance Methods to Enhance the Single-Trial Detection of Event-Related Potentials for a BCI Spelling System | Saavedra C, Salas R, Bougrain L. | Comput Intell Neurosci. 2019 Aug 26;2019:8432953. doi: 10.1155/2019/8432953. eCollection 2019. | Saavedra C | Comput Intell Neurosci | 2019 | 2019-09-20 | PMC6732649 |  | 10.1155/2019/8432953
Mur19 | 194 | 31109062 | An Unsupervised Method for Artefact Removal in EEG Signals | Mur A, Dormido R, Duro N. | Sensors (Basel). 2019 May 18;19(10):2302. doi: 10.3390/s19102302. | Mur A | Sensors (Basel) | 2019 | 2019-05-22 | PMC6567218 |  | 10.3390/s19102302
Blum19 | 196 | 31105543 | A Riemannian Modification of Artifact Subspace Reconstruction for EEG Artifact Handling | Blum S, Jacobsen NSJ, Bleichner MG, Debener S. | Front Hum Neurosci. 2019 Apr 26;13:141. doi: 10.3389/fnhum.2019.00141. eCollection 2019. | Blum S | Front Hum Neurosci | 2019 | 2019-05-21 | PMC6499032 |  | 10.3389/fnhum.2019.00141
Nordin19 | 207 | 30886202 | Human electrocortical dynamics while stepping over obstacles | Nordin AD, Hairston WD, Ferris DP. | Sci Rep. 2019 Mar 18;9(1):4693. doi: 10.1038/s41598-019-41131-2. | Nordin AD | Sci Rep | 2019 | 2019-03-20 | PMC6423113 |  | 10.1038/s41598-019-41131-2
Grosselin19 | 216 | 30709004 | Quality Assessment of Single-Channel EEG for Wearable Devices | Grosselin F, Navarro-Sune X, Vozzi A, Pandremmenou K, De Vico Fallani F, Attal Y, Chavez M. | Sensors (Basel). 2019 Jan 31;19(3):601. doi: 10.3390/s19030601. | Grosselin F | Sensors (Basel) | 2019 | 2019-02-03 | PMC6387437 |  | 10.3390/s19030601
Li18 | 228 | 30618572 | An Approach for Brain-Controlled Prostheses Based on a Facial Expression Paradigm | Li R, Zhang X, Lu Z, Liu C, Li H, Sheng W, Odekhe R. | Front Neurosci. 2018 Dec 18;12:943. doi: 10.3389/fnins.2018.00943. eCollection 2018. | Li R | Front Neurosci | 2018 | 2019-01-09 | PMC6305548 |  | 10.3389/fnins.2018.00943
Chu18 | 240 | 30323737 | A Decoding Scheme for Incomplete Motor Imagery EEG With Deep Belief Network | Chu Y, Zhao X, Zou Y, Xu W, Han J, Zhao Y. | Front Neurosci. 2018 Sep 28;12:680. doi: 10.3389/fnins.2018.00680. eCollection 2018. | Chu Y | Front Neurosci | 2018 | 2018-10-17 | PMC6172343 |  | 10.3389/fnins.2018.00680
San-Martin18 | 241 | 30319337 | A Method for Simultaneous Evaluation of Muscular and Neural Prepulse Inhibition | San-Martin R, Zimiani MI, Noya C, Ávila MAV, Shuhama R, Del-Ben CM, Menezes PR, Fraga FJ, Salum C. | Front Neurosci. 2018 Sep 26;12:654. doi: 10.3389/fnins.2018.00654. eCollection 2018. | San-Martin R | Front Neurosci | 2018 | 2018-10-16 | PMC6168667 |  | 10.3389/fnins.2018.00654
Sebek18 | 248 | 30106969 | Suppression of overlearning in independent component analysis used for removal of muscular artifacts from electroencephalographic records | Sebek J, Bortel R, Sovka P. | PLoS One. 2018 Aug 14;13(8):e0201900. doi: 10.1371/journal.pone.0201900. eCollection 2018. | Sebek J | PLoS One | 2018 | 2018-08-15 | PMC6091961 |  | 10.1371/journal.pone.0201900
Arad18 | 259 | 29768471 | Performance-based approach for movement artifact removal from electroencephalographic data recorded during locomotion | Arad E, Bartsch RP, Kantelhardt JW, Plotnik M. | PLoS One. 2018 May 16;13(5):e0197153. doi: 10.1371/journal.pone.0197153. eCollection 2018. | Arad E | PLoS One | 2018 | 2018-05-17 | PMC5955591 |  | 10.1371/journal.pone.0197153
Symeonidou18 | 265 | 29614020 | Effects of Cable Sway, Electrode Surface Area, and Electrode Mass on Electroencephalography Signal Quality during Motion | Symeonidou ER, Nordin AD, Hairston WD, Ferris DP. | Sensors (Basel). 2018 Apr 3;18(4):1073. doi: 10.3390/s18041073. | Symeonidou ER | Sensors (Basel) | 2018 | 2018-04-04 | PMC5948545 |  | 10.3390/s18041073
Stone18 | 268 | 29618975 | Automatic Removal of Physiological Artifacts in EEG: The Optimized Fingerprint Method for Sports Science Applications | Stone DB, Tamburro G, Fiedler P, Haueisen J, Comani S. | Front Hum Neurosci. 2018 Mar 21;12:96. doi: 10.3389/fnhum.2018.00096. eCollection 2018. | Stone DB | Front Hum Neurosci | 2018 | 2018-04-06 | PMC5871683 |  | 10.3389/fnhum.2018.00096
Frolich18 | 272 | 29322469 | Removal of muscular artifacts in EEG signals: a comparison of linear decomposition methods | Frølich L, Dowding I. | Brain Inform. 2018 Mar;5(1):13-22. doi: 10.1007/s40708-017-0074-6. Epub 2018 Jan 10. | Frølich L | Brain Inform | 2018 | 2018-01-12 | PMC5893498 |  | 10.1007/s40708-017-0074-6
Gabard-Durnam18 | 273 | 29535597 | The Harvard Automated Processing Pipeline for Electroencephalography (HAPPE): Standardized Processing Software for Developmental and High-Artifact Data | Gabard-Durnam LJ, Mendez Leal AS, Wilkinson CL, Levin AR. | Front Neurosci. 2018 Feb 27;12:97. doi: 10.3389/fnins.2018.00097. eCollection 2018. | Gabard-Durnam LJ | Front Neurosci | 2018 | 2018-03-15 | PMC5835235 |  | 10.3389/fnins.2018.00097
Tamburro18 | 275 | 29492336 | A new ICA-based fingerprint method for the automatic removal of physiological artifacts from EEG recordings | Tamburro G, Fiedler P, Stone D, Haueisen J, Comani S. | PeerJ. 2018 Feb 23;6:e4380. doi: 10.7717/peerj.4380. eCollection 2018. | Tamburro G | PeerJ | 2018 | 2018-03-02 | PMC5826009 |  | 10.7717/peerj.4380
Lin18 | 281 | 29599950 | Real-Time EEG Signal Enhancement Using Canonical Correlation Analysis and Gaussian Mixture Clustering | Lin CT, Huang CS, Yang WY, Singh AK, Chuang CH, Wang YK. | J Healthc Eng. 2018 Jan 15;2018:5081258. doi: 10.1155/2018/5081258. eCollection 2018. | Lin CT | J Healthc Eng | 2018 | 2018-03-31 | PMC5823426 |  | 10.1155/2018/5081258
Delisle-Rodriguez17 | 293 | 29186848 | Adaptive Spatial Filter Based on Similarity Indices to Preserve the Neural Information on EEG Signals during On-Line Processing | Delisle-Rodriguez D, Villa-Parra AC, Bastos-Filho T, López-Delis A, Frizera-Neto A, Krishnan S, Rocon E. | Sensors (Basel). 2017 Nov 25;17(12):2725. doi: 10.3390/s17122725. | Delisle-Rodriguez D | Sensors (Basel) | 2017 | 2017-12-01 | PMC5751387 |  | 10.3390/s17122725
Wang17 | 300 | 28714587 | Stretchable Dry Electrodes with Concentric Ring Geometry for Enhancing Spatial Resolution in Electrophysiology | Wang K, Parekh U, Pailla T, Garudadri H, Gilja V, Ng TN. | Adv Healthc Mater. 2017 Oct;6(19). doi: 10.1002/adhm.201700552. Epub 2017 Jul 17. | Wang K | Adv Healthc Mater | 2017 | 2017-07-18 |  |  | 10.1002/adhm.201700552
Jas17 | 301 | 28645840 | Autoreject: Automated artifact rejection for MEG and EEG data | Jas M, Engemann DA, Bekhti Y, Raimondo F, Gramfort A. | Neuroimage. 2017 Oct 1;159:417-429. doi: 10.1016/j.neuroimage.2017.06.030. Epub 2017 Jun 20. | Jas M | Neuroimage | 2017 | 2017-06-25 | PMC7243972 | NIHMS1578244 | 10.1016/j.neuroimage.2017.06.030
Oliveira17 | 328 | 28491016 | A Channel Rejection Method for Attenuating Motion-Related Artifacts in EEG Recordings during Walking | Oliveira AS, Schlink BR, Hairston WD, König P, Ferris DP. | Front Neurosci. 2017 Apr 26;11:225. doi: 10.3389/fnins.2017.00225. eCollection 2017. | Oliveira AS | Front Neurosci | 2017 | 2017-05-12 | PMC5405125 |  | 10.3389/fnins.2017.00225
DelPozo-Banos17 | 333 | 28112387 | Localized component filtering for electroencephalogram artifact rejection | DelPozo-Baños M, Weidemann CT. | Psychophysiology. 2017 Apr;54(4):608-619. doi: 10.1111/psyp.12810. Epub 2017 Jan 23. | DelPozo-Baños M | Psychophysiology | 2017 | 2017-01-24 |  |  | 10.1111/psyp.12810
Weiss17 | 345 | 28491280 | AR2, a novel automatic muscle artifact reduction software method for ictal EEG interpretation: Validation and comparison of performance with commercially available software | Weiss SA, Asadi-Pooya AA, Vangala S, Moy S, Wyeth DH, Orosz I, Gibbs M, Schrader L, Lerner J, Cheng CK, Chang E, Rajaraman R, Keselman I, Churchman P, Bower-Baca C, Numis AL, Ho MG, Rao L, Bhat A, Suski J, Asadollahi M, Ambrose T, Fernandez A, Nei M, Skidmore C, Mintzer S, Eliashiv DS, Mathern GW, Nuwer MR, Sperling M, Engel J Jr, Stern JM. | F1000Res. 2017 Jan 10;6:30. doi: 10.12688/f1000research.10569.2. eCollection 2017. | Weiss SA | F1000Res | 2017 | 2017-05-16 | PMC5399961 |  | 10.12688/f1000research.10569.2
Roy17 | 347 | 29118966 | Gaussian Elimination-Based Novel Canonical Correlation Analysis Method for EEG Motion Artifact Removal | Roy V, Shukla S, Shukla PK, Rawat P. | J Healthc Eng. 2017;2017:9674712. doi: 10.1155/2017/9674712. Epub 2017 Oct 8. | Roy V | J Healthc Eng | 2017 | 2017-11-10 | PMC5651166 |  | 10.1155/2017/9674712
Melman16 | 370 | 27102041 | Robust power spectral estimation for EEG data | Melman T, Victor JD. | J Neurosci Methods. 2016 Aug 1;268:14-22. doi: 10.1016/j.jneumeth.2016.04.015. Epub 2016 Apr 19. | Melman T | J Neurosci Methods | 2016 | 2016-04-23 | PMC4903894 | NIHMS785163 | 10.1016/j.jneumeth.2016.04.015
Oliveira16 | 377 | 27242467 | Proposing Metrics for Benchmarking Novel EEG Technologies Towards Real-World Measurements | Oliveira AS, Schlink BR, Hairston WD, König P, Ferris DP. | Front Hum Neurosci. 2016 May 10;10:188. doi: 10.3389/fnhum.2016.00188. eCollection 2016. | Oliveira AS | Front Hum Neurosci | 2016 | 2016-06-01 | PMC4861738 |  | 10.3389/fnhum.2016.00188
Cheveigne16 | 389 | 26778608 | Sparse time artifact removal | de Cheveigné A. | J Neurosci Methods. 2016 Mar 15;262:14-20. doi: 10.1016/j.jneumeth.2016.01.005. Epub 2016 Jan 8. | de Cheveigné A | J Neurosci Methods | 2016 | 2016-01-19 |  |  | 10.1016/j.jneumeth.2016.01.005
Costa16 | 392 | 26941601 | Characterization of Artifacts Produced by Gel Displacement on Non-invasive Brain-Machine Interfaces during Ambulation | Costa Á, Salazar-Varas R, Úbeda A, Azorín JM. | Front Neurosci. 2016 Feb 25;10:60. doi: 10.3389/fnins.2016.00060. eCollection 2016. | Costa Á | Front Neurosci | 2016 | 2016-03-05 | PMC4766300 |  | 10.3389/fnins.2016.00060
Nathan16 | 400 | 26793089 | Negligible Motion Artifacts in Scalp Electroencephalography (EEG) During Treadmill Walking | Nathan K, Contreras-Vidal JL. | Front Hum Neurosci. 2016 Jan 13;9:708. doi: 10.3389/fnhum.2015.00708. eCollection 2015. | Nathan K | Front Hum Neurosci | 2016 | 2016-01-22 | PMC4710850 |  | 10.3389/fnhum.2015.00708
Zou16 | 405 | 25415992 | Automatic Identification of Artifact-Related Independent Components for Artifact Removal in EEG Recordings | Zou Y, Nathan V, Jafari R. | IEEE J Biomed Health Inform. 2016 Jan;20(1):73-81. doi: 10.1109/JBHI.2014.2370646. Epub 2014 Nov 13. | Zou Y | IEEE J Biomed Health Inform | 2016 | 2014-11-22 | PMC5309922 | NIHMS749533 | 10.1109/JBHI.2014.2370646
Gordon15 | 406 | 26306657 | Informed decomposition of electroencephalographic data | Gordon SM, Lawhern V, Passaro AD, McDowell K. | J Neurosci Methods. 2015 Dec 30;256:41-55. doi: 10.1016/j.jneumeth.2015.08.019. Epub 2015 Aug 22. | Gordon SM | J Neurosci Methods | 2015 | 2015-08-27 |  |  | 10.1016/j.jneumeth.2015.08.019
Snyder15 | 408 | 26648858 | Independent Component Analysis of Gait-Related Movement Artifact Recorded using EEG Electrodes during Treadmill Walking | Snyder KL, Kline JE, Huang HJ, Ferris DP. | Front Hum Neurosci. 2015 Dec 1;9:639. doi: 10.3389/fnhum.2015.00639. eCollection 2015. | Snyder KL | Front Hum Neurosci | 2015 | 2015-12-10 | PMC4664645 |  | 10.3389/fnhum.2015.00639
Kline15 | 420 | 26083595 | Isolating gait-related movement artifacts in electroencephalography during human walking | Kline JE, Huang HJ, Snyder KL, Ferris DP. | J Neural Eng. 2015 Aug;12(4):046022. doi: 10.1088/1741-2560/12/4/046022. Epub 2015 Jun 17. | Kline JE | J Neural Eng | 2015 | 2015-06-18 | PMC4946867 | NIHMS702844 | 10.1088/1741-2560/12/4/046022
Raduntz15 | 431 | 25666892 | EEG artifact elimination by extraction of ICA-component features using image processing algorithms | Radüntz T, Scouten J, Hochmuth O, Meffert B. | J Neurosci Methods. 2015 Mar 30;243:84-93. doi: 10.1016/j.jneumeth.2015.01.030. Epub 2015 Feb 7. | Radüntz T | J Neurosci Methods | 2015 | 2015-02-11 |  |  | 10.1016/j.jneumeth.2015.01.030
Mariani15 | 444 | 26738116 | Use of multiscale entropy to facilitate artifact detection in electroencephalographic signals | Mariani S, Borges AF, Henriques T, Goldberger AL, Costa MD. | Annu Int Conf IEEE Eng Med Biol Soc. 2015;2015:7869-72. doi: 10.1109/EMBC.2015.7320216. | Mariani S | Annu Int Conf IEEE Eng Med Biol Soc | 2015 | 2016-01-07 | PMC4855284 | NIHMS779763 | 10.1109/EMBC.2015.7320216
Zhang15 | 447 | 26380294 | Automatic Artifact Removal from Electroencephalogram Data Based on A Priori Artifact Information | Zhang C, Tong L, Zeng Y, Jiang J, Bu H, Yan B, Li J. | Biomed Res Int. 2015;2015:720450. doi: 10.1155/2015/720450. Epub 2015 Aug 25. | Zhang C | Biomed Res Int | 2015 | 2015-09-18 | PMC4562337 |  | 10.1155/2015/720450
Bulea14 | 451 | 25505377 | Sitting and standing intention can be decoded from scalp EEG recorded prior to movement execution | Bulea TC, Prasad S, Kilicarslan A, Contreras-Vidal JL. | Front Neurosci. 2014 Nov 25;8:376. doi: 10.3389/fnins.2014.00376. eCollection 2014. | Bulea TC | Front Neurosci | 2014 | 2014-12-16 | PMC4243562 |  | 10.3389/fnins.2014.00376
Chen14 | 455 | 25275348 | A preliminary study of muscular artifact cancellation in single-channel EEG | Chen X, Liu A, Peng H, Ward RK. | Sensors (Basel). 2014 Oct 1;14(10):18370-89. doi: 10.3390/s141018370. | Chen X | Sensors (Basel) | 2014 | 2014-10-03 | PMC4239950 |  | 10.3390/s141018370
Yilmaz14 | 465 | 25071531 | Interference of tonic muscle activity on the EEG: a single motor unit study | Yilmaz G, Ungan P, Sebik O, Uginčius P, Türker KS. | Front Hum Neurosci. 2014 Jul 11;8:504. doi: 10.3389/fnhum.2014.00504. eCollection 2014. | Yilmaz G | Front Hum Neurosci | 2014 | 2014-07-30 | PMC4092367 |  | 10.3389/fnhum.2014.00504
Reis14 | 482 | 24715858 | Methodological aspects of EEG and body dynamics measurements during motion | Reis PM, Hebenstreit F, Gabsteiger F, von Tscharner V, Lochmann M. | Front Hum Neurosci. 2014 Mar 24;8:156. doi: 10.3389/fnhum.2014.00156. eCollection 2014. | Reis PM | Front Hum Neurosci | 2014 | 2014-04-10 | PMC3970018 |  | 10.3389/fnhum.2014.00156
Boudet14 | 490 | 25298967 | Dual adaptive filtering by optimal projection applied to filter muscle artifacts on EEG and comparative study | Boudet S, Peyrodie L, Szurhaj W, Bolo N, Pinti A, Gallois P. | ScientificWorldJournal. 2014;2014:374679. doi: 10.1155/2014/374679. Epub 2014 Sep 14. | Boudet S | ScientificWorldJournal | 2014 | 2014-10-10 | PMC4178918 |  | 10.1155/2014/374679
Li13 | 509 | 23919646 | L1 norm based common spatial patterns decomposition for scalp EEG BCI | Li P, Xu P, Zhang R, Guo L, Yao D. | Biomed Eng Online. 2013 Aug 6;12:77. doi: 10.1186/1475-925X-12-77. | Li P | Biomed Eng Online | 2013 | 2013-08-08 | PMC3750597 |  | 10.1186/1475-925X-12-77
Lawhern13 | 524 | 23638169 | DETECT: a MATLAB toolbox for event detection and identification in time series, with applications to artifact detection in EEG signals | Lawhern V, Hairston WD, Robbins K. | PLoS One. 2013 Apr 24;8(4):e62944. doi: 10.1371/journal.pone.0062944. Print 2013. | Lawhern V | PLoS One | 2013 | 2013-05-03 | PMC3634773 |  | 10.1371/journal.pone.0062944
Nottage13 | 538 | 22965826 | A novel method for reducing the effect of tonic muscle activity on the gamma band of the scalp EEG | Nottage JF, Morrison PD, Williams SC, Ffytche DH. | Brain Topogr. 2013 Jan;26(1):50-61. doi: 10.1007/s10548-012-0255-9. Epub 2012 Sep 11. | Nottage JF | Brain Topogr | 2013 | 2012-09-12 | PMC3536945 |  | 10.1007/s10548-012-0255-9
Yong12 | 549 | 22838499 | Automatic artefact removal in a self-paced hybrid brain- computer interface system | Yong X, Fatourechi M, Ward RK, Birch GE. | J Neuroeng Rehabil. 2012 Jul 27;9:50. doi: 10.1186/1743-0003-9-50. | Yong X | J Neuroeng Rehabil | 2012 | 2012-07-31 | PMC3480939 |  | 10.1186/1743-0003-9-50
Lau12 | 550 | 22828128 | Weighted phase lag index stability as an artifact resistant measure to detect cognitive EEG activity during locomotion | Lau TM, Gwin JT, McDowell KG, Ferris DP. | J Neuroeng Rehabil. 2012 Jul 24;9:47. doi: 10.1186/1743-0003-9-47. | Lau TM | J Neuroeng Rehabil | 2012 | 2012-07-26 | PMC3488562 |  | 10.1186/1743-0003-9-47
Lawhern12 | 551 | 22634706 | Detection and classification of subject-generated artifacts in EEG signals using autoregressive models | Lawhern V, Hairston WD, McDowell K, Westerfield M, Robbins K. | J Neurosci Methods. 2012 Jul 15;208(2):181-9. doi: 10.1016/j.jneumeth.2012.05.017. Epub 2012 May 23. | Lawhern V | J Neurosci Methods | 2012 | 2012-05-29 |  |  | 10.1016/j.jneumeth.2012.05.017
Winkler11 | 583 | 21810266 | Automatic classification of artifactual ICA-components for artifact removal in EEG signals | Winkler I, Haufe S, Tangermann M. | Behav Brain Funct. 2011 Aug 2;7:30. doi: 10.1186/1744-9081-7-30. | Winkler I | Behav Brain Funct | 2011 | 2011-08-04 | PMC3175453 |  | 10.1186/1744-9081-7-30
Gwin10 | 615 | 20410364 | Removal of movement artifact from high-density EEG recorded during walking and running | Gwin JT, Gramann K, Makeig S, Ferris DP. | J Neurophysiol. 2010 Jun;103(6):3526-34. doi: 10.1152/jn.00105.2010. Epub 2010 Apr 21. | Gwin JT | J Neurophysiol | 2010 | 2010-04-23 | PMC3774587 |  | 10.1152/jn.00105.2010
McMenamin10 | 624 | 19833218 | Validation of ICA-based myogenic artifact correction for scalp and source-localized EEG | McMenamin BW, Shackman AJ, Maxwell JS, Bachhuber DR, Koppenhaver AM, Greischar LL, Davidson RJ. | Neuroimage. 2010 Feb 1;49(3):2416-32. doi: 10.1016/j.neuroimage.2009.10.010. Epub 2009 Oct 13. | McMenamin BW | Neuroimage | 2010 | 2009-10-17 | PMC2818255 | NIHMS150978 | 10.1016/j.neuroimage.2009.10.010
Mosher09 | 640 | 20234848 | GENERALIZED SIDELOBE CANCELLER FOR MAGNETOENCEPHALOGRAPHY ARRAYS | Mosher JC, Hämäläinen MS, Pantazis D, Hui HB, Burgess RC, Leahy RM. | Proc IEEE Int Symp Biomed Imaging. 2009 Aug 7;2009:149-152. doi: 10.1109/ISBI.2009.5193005. | Mosher JC | Proc IEEE Int Symp Biomed Imaging | 2009 | 2010-03-18 | PMC2838207 | NIHMS164379 | 10.1109/ISBI.2009.5193005
McMenamin09 | 648 | 19298626 | Validation of regression-based myogenic correction techniques for scalp and source-localized EEG | McMenamin BW, Shackman AJ, Maxwell JS, Greischar LL, Davidson RJ. | Psychophysiology. 2009 May;46(3):578-92. doi: 10.1111/j.1469-8986.2009.00787.x. Epub 2009 Mar 4. | McMenamin BW | Psychophysiology | 2009 | 2009-03-21 | PMC2677703 | NIHMS82796 | 10.1111/j.1469-8986.2009.00787.x
Delorme04 | 707 | 15102499 | EEGLAB: an open source toolbox for analysis of single-trial EEG dynamics including independent component analysis | Delorme A, Makeig S. | J Neurosci Methods. 2004 Mar 15;134(1):9-21. doi: 10.1016/j.jneumeth.2003.10.009. | Delorme A | J Neurosci Methods | 2004 | 2004-04-23 |  |  | 10.1016/j.jneumeth.2003.10.009
Iriarte03 | 713 | 14530738 | Independent component analysis as a tool to eliminate artifacts in EEG: a quantitative study | Iriarte J, Urrestarazu E, Valencia M, Alegre M, Malanda A, Viteri C, Artieda J. | J Clin Neurophysiol. 2003 Jul-Aug;20(4):249-57. doi: 10.1097/00004691-200307000-00004. | Iriarte J | J Clin Neurophysiol | 2003 | 2003-10-08 |  |  | 10.1097/00004691-200307000-00004
Jung01 | 725 | 20824156 | Imaging Brain Dynamics Using Independent Component Analysis | Jung TP, Makeig S, McKeown MJ, Bell AJ, Lee TW, Sejnowski TJ. | Proc IEEE Inst Electr Electron Eng. 2001 Jul 1;89(7):1107-1122. doi: 10.1109/5.939827. | Jung TP | Proc IEEE Inst Electr Electron Eng | 2001 | 2010-09-09 | PMC2932458 | HHMIMS197143 | 10.1109/5.939827

## Filter rules
Index | Reason
0 | NO_OPEN_ACCESS
1 | IS_SYSTEMATIC_REVIEW
2 | NO_NI_HUMAN_EEG
3 | NO_ARTIFACT_FOCUS
4 | NO_MOTION_FOCUS
5 | NO_REAL_DATASET
6 | NO_SCIENCE_GRADE

## FitOnline Score
Fit score
0 | means | has never been demonstrated to work at all with any real data
requires preconditions that can't be met by online scenarios at all (e.g. req. fNIRS setup)
all or most trials/epochs/windows contaminated with complex motion artifacts were removed
is not generalizable to the intended target population
This is a default pipeline with no motion artifact filter, e.g., Original EEG
1 | means | default if 0 conditions and 2 conditions don't apply
2 | means | proposed algorithm/method/pipeline is already in usage somewhere i real aplications or similar
proposed algorithm/method/pipeline corrects motion artifacts without removing entire time windows
Correction score
0 | means | does not show how a pipeline detects or corrects myogenic artifacts, but may provide ways to mitigate motion artifacts from e.g. cable swings
the pipeline is not tested how it reduces myogenic artifacts
1 | means | shows how the pipeline detects artifacts in and that the results enhances by removing them
2 | means | shows how the pipeline (detects and) corrects artifacts from the EEG such that the EEG can be provided with reduced artifacts over the whole duration
Online Score
0 | means | has shown to be very slow/imperformant (multiple seconds delay). E.g. if the signal is filtered without or within very wide windows.
is not shown to be functional in online scenarios (e.g. setting is not transferable to our scenario controlling devices in RT)
requires preconditions that can't be met by online scenarios at all (e.g. req. fNIRS setup)
1 | means | default if 0 conditions and 2 conditions don't apply
2 | means | has been validated with online RT communication (e.g. sending command to device)
the pipeline has been shown to work in online studies
this scala will be applied to all newly introduced methods/algo/pipeline

## Mental Strategy
name | description | examples | details
brain-teaser tasks | tasks that require the user to solve a specific mental problem, such as word association and mental subtraction
motor imagery tasks | imagine movement of e.g. left/right hand or feet, etc. | MR task, SMR (sensori motor rythm)
non-motor imagery tasks | auditory imagery and spatial navigation
dynamic visualization tasks | mental rotation of e.g., an imagined object
attention strategy | SSVEP (steady state visually evoked potential), SSAEP (steady state auditory evoked potential), SSSTEP (steady state somatosensory evoked potential), mVEP (motion-onset visual evoked potential) | SSVEP, SSAEP, SSSEP, mVEP, exposure to visual or auditive stimuli | Electrical response on external stimulus, as positive flanks after a characteristic time lapse such as (P50, N100, P200, N200, P300, N300, P400, N400, P600, mismatch negativity, contingent negative variation, and movement-related cortical potentials)
motor execution | execution of movement
- | No strategy is used, measuring the EEG without mental tasks or external stimuli

## Authors
Reference | Author (alle Autoren angeben) | Institution
Hossain22 | Hossain MS | University Kebangsaan Malaysia, Malaysia
Hossain22 | Chowdhury MEH | Qatar University, Qatar
Hossain22 | Reaz MBI | University Kebangsaan Malaysia, Malaysia
Hossain22 | Ali SHM | University Kebangsaan Malaysia, Malaysia
Hossain22 | Bakar AAA | University Kebangsaan Malaysia, Malaysia
Hossain22 | Kiranyaz S | Qatar University, Qatar
Hossain22 | Khandakar A | Qatar University, Qatar
Hossain22 | Alhatou M | Hamad General Hospital, Qatar
Hossain22 | Habib R | BIRDEM General Hospital, Bangladesh
Hossain22 | Hossain MM | North South University, Bangladesh
Abu-Farha22 | Abu Farha N | American University of Sharjah, United Arab Emirates
Abu-Farha22 | Al-Shargie F | American University of Sharjah, United Arab Emirates
Abu-Farha22 | Tariq U | American University of Sharjah, United Arab Emirates
Abu-Farha22 | Al-Nashash H | American University of Sharjah, United Arab Emirates
Flo22 | Flo A | Universite Paris-Saclay, France
Flo22 | Gennari G | Universite Paris-Saclay, France
Flo22 | Benjamin L | Universite Paris-Saclay, France
Flo22 | Dehaene-Lambertz G | Universite Paris-Saclay, France
Kumaravel22 | Kumaravel VP | Fondazione Bruno Kessler, Italy
Kumaravel22 | Kumaravel VP | University of Trento, Italy
Kumaravel22 | Farella E | Fondazione Bruno Kessler, Italy
Kumaravel22 | Parise E | University of Trento, Italy
Kumaravel22 | Buiatti M | University of Trento, Italy
Soghoyan21 | Soghoyan G | National Research University Higher School of Economics, Russia
Soghoyan21 | Soghoyan G | Russian Academy of Sciences, Russia
Soghoyan21 | Ledovsky A | Russian Academy of Sciences, Russia
Soghoyan21 | Ledovsky A | Skolkovo Institute of Science and Technology, Russia
Soghoyan21 | Nekrashevich M | Skolkovo Institute of Science and Technology, Russia
Soghoyan21 | Martynova O | Russian Academy of Sciences, Russia
Soghoyan21 | Polikanova I | National Research University Higher School of Economics, Russia
Soghoyan21 | Portnova G | Russian Academy of Sciences, Russia
Soghoyan21 | Rebreikina A | Russian Academy of Sciences, Russia
Soghoyan21 | Sysoeva O | Russian Academy of Sciences, Russia
Soghoyan21 | Sharaev M | Russian Academy of Sciences, Russia
Soghoyan21 | Sharaev M | Skolkovo Institute of Science and Technology, Russia
Karpiel21 | Karpiel I | Lukasiewicz Research Network, Poland
Karpiel21 | Kurasz Z | University of Silesia, Poland
Karpiel21 | Kurasz R | Independent, Poland
Karpiel21 | Duch K | University of Silesia, Poland
Beach21 | Beach C | The University of Manchester, UK
Beach21 | Li M | Arm, UK
Beach21 | Balaban E | The University of Manchester, UK
Beach21 | Casson AJ | The University of Manchester, UK
Dai21 | Dai Y | Nankai University, China
Dai21 | Duan F | Nankai University, China
Dai21 | Feng F | Nankai University, China
Dai21 | Sun Z | RIKEN, Japan
Dai21 | Zhang Y | Lehigh University Bethlehem, USA
Dai21 | Ciaafa CF | Nankai University, China
Dai21 | Ciaafa CF | Instituto Argentino De Radioastronomia, Argentina
Dai21 | Marti-Puig P | Central University of Catalonia, Spain
Dai21 | Sole-Casals J | Nankai University, China
Dai21 | Sole-Casals J | Central University of Catalonia, Spain
Tost21 | Tost A | Universitat Politecnica de Catalunya, Spain
Tost21 | Migliorelli C | Universitat Politecnica de Catalunya, Spain
Tost21 | Migliorelli C | CIBER-BBN, Spain
Tost21 | Migliorelli C | Institute de Recerca Sant Joan de Deu, Spain
Tost21 | Bachiller A | Universitat Politecnica de Catalunya, Spain
Tost21 | Bachiller A | CIBER-BBN, Spain
Tost21 | Bachiller A | Institute de Recerca Sant Joan de Deu, Spain
Tost21 | Medina-Rivera I | Institute de Recerca Sant Joan de Deu, Spain
Tost21 | Romero S | Universitat Politecnica de Catalunya, Spain
Tost21 | Romero S | CIBER-BBN, Spain
Tost21 | Romero S | Institute de Recerca Sant Joan de Deu, Spain
Tost21 | Garcia-Cazorla A | Institute de Recerca Sant Joan de Deu, Spain
Tost21 | Garcia-Cazorla A | Hospital Sant Joan De Deu, Spain
Tost21 | Mananas MA | Universitat Politecnica de Catalunya, Spain
Tost21 | Mananas MA | CIBER-BBN, Spain
Tost21 | Mananas MA | Institute de Recerca Sant Joan de Deu, Spain
Kaur21 | Kaur C | Panjab University Chandigarh, India
Kaur21 | Singh P | Panjab University Chandigarh, India
Kaur21 | Sahni S | Cheema Medical Complex, India
Rosanne21 | Rosanne O | Université du Québec, Canada
Rosanne21 | Albuquerque I | Université du Québec, Canada
Rosanne21 | Cassani R | Université du Québec, Canada
Rosanne21 | Gagnon JF | Quebec City, Canada
Rosanne21 | Tremblay S | Université Laval, Canada
Rosanne21 | Falk T | Université du Québec, Canada
Nahmias21 | Nahmias DO | Center for Devices and Radiological Health (Silver Spring, Maryland), USA
Nahmias21 | Nahmias DO | University of Maryland, USA
Nahmias21 | Kontson KL | Center for Devices and Radiological Health (Silver Spring, Maryland), USA
Saba-Sadiya21 | Saba-Sadiya S | Michigan State University, USA
Saba-Sadiya21 | Saba-Sadiya S | Michigan State University, USA
Saba-Sadiya21 | Chantland E | Michigan State University, USA
Saba-Sadiya21 | Alhanai T | New York University Abu Dhabi, United Arab Emirates
Saba-Sadiya21 | Liu T | Michigan State University, USA
Saba-Sadiya21 | Ghassemi MM | Michigan State University, USA
Li21 | Li Y | University of California, USA
Li21 | Wang PT | University of California, USA
Li21 | Vaidya MP | Northwestern University Chicago, USA
Li21 | Flint RD | Northwestern University Chicago, USA
Li21 | Liu CY | University of Southern California, USA
Li21 | Liu CY | Rancho Los Amigos National Rehabilitation Center, USA
Li21 | Slutzky MW | Northwestern University Chicago, USA
Li21 | Do AH | University of California, USA
Anders20 | Anders P | Norwegian University of Science and Technology, Norway
Anders20 | Müller H | Paderborn University, Germany
Anders20 | Skjaeret-Maroni N | Norwegian University of Science and Technology, Norway
Anders20 | Vereijken B | Norwegian University of Science and Technology, Norway
Anders20 | Baumeister J | Paderborn University, Germany
Zhang20 | Zhang L | National University of Singapore, Singapore
Zhang20 | Kumar KS | National University of Singapore, Singapore
Zhang20 | He H | National University of Singapore, Singapore
Zhang20 | Cai CJ | National University of Singapore, Singapore
Zhang20 | Cai CJ | Singapore Institute of Manufacturing Technology, Singapore
Zhang20 | He X | National University of Singapore, Singapore
Zhang20 | Gao H | National University of Singapore, Singapore
Zhang20 | Gao H | National University of Singapore Research Institute, China
Zhang20 | Yue S | National University of Singapore, Singapore
Zhang20 | Li C | National University of Singapore, Singapore
Zhang20 | Li C | National University of Singapore Research Institute, China
Zhang20 | Li C | Beijing Advanced Innovation Center for Intelligent Robots and Systems, China
Zhang20 | Seet RCS | National University Health System, Singapore
Zhang20 | Seet RCS | National University of Singapore, Singapore
Zhang20 | Ren H | National University of Singapore, Singapore
Zhang20 | Ren H | National University of Singapore Research Institute, China
Zhang20 | Ren H | The Chinese University of Hong Kong Robotics Institute, China
Zhang20 | Ouyang J | National University of Singapore, Singapore
Leach20 | Leach SC | University of Maryland, USA
Leach20 | Morales S | University of Maryland, USA
Leach20 | Bowers ME | University of Maryland, USA
Leach20 | Buzzell GA | University of Maryland, USA
Leach20 | Debnath R | University of Maryland, USA
Leach20 | Beall D | University of Maryland, USA
Leach20 | Fox NA | University of Maryland, USA
Saini20 | Saini M | Shiv Nadar University, India
Saini20 | Satija U | Indian Institute of Technology Patna, India
Saini20 | Upadhayay MD | Shiv Nadar University, India
Mucarquer20 | Mucarquer JA | Universidad Técnica Federico Santa María, Chile
Mucarquer20 | Prado P | Universidad Técnica Federico Santa María, Chile
Mucarquer20 | Escobar MJ | Universidad Técnica Federico Santa María, Chile
Mucarquer20 | El-Deredy W | Universidad de Valparaiso, Chile
Mucarquer20 | Zañartu M | Universidad Técnica Federico Santa María, Chile
Nordin20 | Nordin AD | University of Florida, USA
Nordin20 | Hairston WD | US Army Research Laboratory, USA
Nordin20 | Ferris DP | University of Florida, USA
Liu19 | Liu Q | University of Science and Technology of China, China
Liu19 | Liu A | University of Science and Technology of China, China
Liu19 | Zhang X | University of Science and Technology of China, China
Liu19 | Chen X | University of Science and Technology of China, China
Liu19 | Qian R | The first Affiliated Hospital of University of Science and Technology of China, China
Liu19 | Chen X | University of Science and Technology of China, China
Benda19 | Benda M | Rhine-Waal University of Applied Sciences, Germany
Benda19 | Volosyak I | Rhine-Waal University of Applied Sciences, Germany
Val-Calvo19 | Val-Calvo M | Universidad Politécnica de Cartagena, Spain
Val-Calvo19 | Val-Calvo M | UNED, Spain
Val-Calvo19 | Álvarez-Sánchez JR | UNED, Spain
Val-Calvo19 | Ferrández-Vicent JM | Universidad Politécnica de Cartagena, Spain
Val-Calvo19 | Fernández E | CIBER-BBN, Spain
Val-Calvo19 | Fernández E | Universidad Miguel Hernández, Alicante, Spain
Saavedra19 | Saavedra C | Universidad de Valparaiso, Chile
Saavedra19 | Saavedra C | Universit´e de Lorraine, France
Saavedra19 | Salas R | Universidad de Valparaiso, Chile
Saavedra19 | Bougrain L | Universit´e de Lorraine, France
Mur19 | Mur A | Universidad Nacional de Educacion a Distancia, Spain
Mur19 | Dormido R | Universidad Nacional de Educacion a Distancia, Spain
Mur19 | Duro N | Universidad Nacional de Educacion a Distancia, Spain
Blum19 | Blum S | University of Oldenburg, Germany
Blum19 | Jacobsen NSJ | University of Oldenburg, Germany
Blum19 | Bleichner MG | University of Oldenburg, Germany
Blum19 | Debener S | University of Oldenburg, Germany
Nordin19 | Nordin AD | University of Florida, USA
Nordin19 | Hairston WD | US Army Research Laboratory, USA
Nordin19 | Ferris DP | University of Florida, USA
Grosselin19 | Grosselin F | Sorbonne Université, France
Grosselin19 | Grosselin F | myBrainTechnologies, France
Grosselin19 | Grosselin F | INRIA, France
Grosselin19 | Navarro-Sune X | myBrainTechnologies, France
Grosselin19 | Vozzi A | myBrainTechnologies, France
Grosselin19 | Pandremmenou K | myBrainTechnologies, France
Grosselin19 | Fallani FDV | Sorbonne Université, France
Grosselin19 | Fallani FDV | INRIA, France
Grosselin19 | Attal Y | myBrainTechnologies, France
Grosselin19 | Chavez M | Hôpital de la Salpêtrière, France
Li18 | Li R | Xi'an Jiaotong University, China
Li18 | Zhang X | Xi'an Jiaotong University, China
Li18 | Lu Z | Xi'an Jiaotong University, China
Li18 | Liu C | Xi'an Jiaotong University, China
Li18 | Li H | Xi'an Jiaotong University, China
Li18 | Sheng W | Oklahome State University, USA
Li18 | Sheng W | Shenzhen Academy of Robotics, China
Li18 | Odekhe R | Xi'an Jiaotong University, China
Chu18 | Chu Y | Chinese Academy of Sciences, China
Chu18 | Chu Y | University of Chinese Academy of Sciences, China
Chu18 | Zhao X | Chinese Academy of Sciences, China
Chu18 | Zou Y | Chinese Academy of Sciences, China
Chu18 | Zou Y | University of Chinese Academy of Sciences, China
Chu18 | Xu W | Chinese Academy of Sciences, China
Chu18 | Xu W | University of Auckland, New Zealand
Chu18 | Han J | Chinese Academy of Sciences, China
Chu18 | Zhao Y | Chinese Academy of Sciences, China
San-Martin18 | San-Martin R | Universidade Federal do ABC, Brazil
San-Martin18 | Zimiani MI | Universidade Federal do ABC, Brazil
San-Martin18 | Noya C | Universidade Federal do ABC, Brazil
San-Martin18 | Avila MAV | Universidade de Sao Paulo, Brazil
San-Martin18 | Shuhama R | Universidade de Sao Paulo, Brazil
San-Martin18 | Del-Ben CM | Universidade de Sao Paulo, Brazil
San-Martin18 | Menezes PR | Universidade de Sao Paulo, Brazil
San-Martin18 | Fraga FJ | Universidade Federal do ABC, Brazil
San-Martin18 | Salum C | Universidade Federal do ABC, Brazil
Sebek18 | Sebek J | Czech Technical University, Czech Republic
Sebek18 | Bortel R | Czech Technical University, Czech Republic
Sebek18 | Sovka P | Czech Technical University, Czech Republic
Arad18 | Arad E | Sheba Medical Center, Israel
Arad18 | Bartsch RP | Bar-Ilan University, Israel
Arad18 | Kantelhardt JW | Martin-Luther-University, Germany
Arad18 | Plotnik M | Sheba Medical Center, Israel
Arad18 | Plotnik M | Tel-Aviv University, Israel
Symeonidou18 | Symeonidou ER | University of Florida, USA
Symeonidou18 | Symeonidou ER | Max Planck Institute for Biological Cybernetics, Germany
Symeonidou18 | Symeonidou ER | International Max Planck Research School for Cognitive and Systems Neuroscience, Germany
Symeonidou18 | Nordin AD | University of Florida, USA
Symeonidou18 | Hairston WD | US Army Research Laboratory, USA
Symeonidou18 | Ferris DP | University of Florida, USA
Stone18 | Stone DB | Università degli Studi G. d’Annunzio Chieti e Pescara, Italy
Stone18 | Tamburro G | Università degli Studi G. d’Annunzio Chieti e Pescara, Italy
Stone18 | Fiedler P | Universität Ilmenau, Germany
Stone18 | Haueisen J | Universität Ilmenau, Germany
Stone18 | Comani S | Università degli Studi G. d’Annunzio Chieti e Pescara, Italy
Frolich18 | Frohlich L | Technical University of Denmark, Denmark
Frolich18 | Dowding I | Technische Universität Berlin, Germany
Gabard-Durnam18 | Gabard-Durnam LJ | Harvard University, USA
Gabard-Durnam18 | Leal ASM | Harvard University, USA
Gabard-Durnam18 | Wilkinson CL | Harvard University, USA
Gabard-Durnam18 | Levin AR | Harvard University, USA
Gabard-Durnam18 | Levin AR | Boston Children’s Hospital, USA
Tamburro18 | Tamburo G | D'Annunzio University of Chieti–Pescara, Italy
Tamburro18 | Fiedler P | Casa di Cura Privata Villa Serena, Italy
Tamburro18 | Fiedler P | Technische Universität Ilmenau, Germany
Tamburro18 | Stone D | D'Annunzio University of Chieti–Pescara, Italy
Tamburro18 | Haueisen J | Technische Universität Ilmenau, Germany
Tamburro18 | Comani S | D'Annunzio University of Chieti–Pescara, Italy
Tamburro18 | Comani S | Casa di Cura Privata Villa Serena, Italy
Lin18 | Lin CT | University of Technology Sydney, Australia
Lin18 | Lin CT | National Chiao Tung University, Taiwan
Lin18 | Huang CS | National Chiao Tung University, Taiwan
Lin18 | Yang WY | National Chiao Tung University, Taiwan
Lin18 | Singh AK | University of Technology Sydney, Australia
Lin18 | Singh AK | National Chiao Tung University, Taiwan
Lin18 | Singh AK | National Chiao Tung University, Taiwan
Lin18 | Chuang CH | University of Technology Sydney, Australia
Lin18 | Chuang CH | National Chiao Tung University, Taiwan
Lin18 | Wang YK | University of Technology Sydney, Australia
Delisle-Rodriguez17 | Delisle-Rodriguez D | Federal University of Espirito Santo, Brazil
Delisle-Rodriguez17 | Delisle-Rodriguez D | University of Oriente, Cuba
Delisle-Rodriguez17 | Villa-Parra AC | Federal University of Espirito Santo, Brazil
Delisle-Rodriguez17 | Villa-Parra AC | Universidad Politécnica Salesiana, Ecuador
Delisle-Rodriguez17 | Bastos-Filho T | Federal University of Espirito Santo, Brazil
Delisle-Rodriguez17 | López-Delis A | University of Oriente, Cuba
Delisle-Rodriguez17 | Frizera-Neto A | Federal University of Espirito Santo, Brazil
Delisle-Rodriguez17 | Krishnan S | Ryerson University, Canada
Delisle-Rodriguez17 | Rocon E | Centre for Automation and Robotics, Spain
Wang17 | Wank K | University of California San Diego, USA
Wang17 | Parekh U | University of California San Diego, USA
Wang17 | Pailla T | University of California San Diego, USA
Wang17 | Garudadri H | University of California San Diego, USA
Wang17 | Gilja V | University of California San Diego, USA
Wang17 | Ng TN | University of California San Diego, USA
Jas17 | Jas M | Université Paris-Saclay, France
Jas17 | Engemann DA | INRIA Saclay - Ile de France, France
Jas17 | Engemann DA | Université Paris-Sacla, France
Jas17 | Engemann DA | ICM PICNIC Lab, France
Jas17 | Bekhti Y | Université Paris-Saclay, France
Jas17 | Raimondo F | ICM PICNIC Lab, France
Jas17 | Raimondo F | Universidad de Buenos Aires, Argentina
Jas17 | Raimondo F | CONICET, Argentina
Jas17 | Raimondo F | Sorbonne Universités, France
Jas17 | Gramfort A | Université Paris-Saclay, France
Oliveira17 | Oliveira AS | University of Michigan, USA
Oliveira17 | Oliveira AS | Aalborg University, Denmark
Oliveira17 | Schlink BR | Aalborg University, Denmark
Oliveira17 | Hairston WD | Aberdeen Proving Ground, USA
Oliveira17 | König P | University of Osnabrück, Germany
Oliveira17 | König P | University Medical Center Hamburg-Eppendorf, Germany
Oliveira17 | Ferris DP | University of Michigan, USA
DelPozo-Banos17 | DelPozo-Banos M | Swansea University, UK
DelPozo-Banos17 | Weidemann CT | Swansea University, UK
Weiss17 | Weiss SA | University of California Los Angeles, USA
Weiss17 | Weiss SA | Thomas Jefferson University, USA
Weiss17 | Asadi-Pooya AA | Thomas Jefferson University, USA
Weiss17 | Vangala S | University of California Los Angeles, USA
Weiss17 | Moy S | University of California Los Angeles, USA
Weiss17 | Wyeth DH | Thomas Jefferson University, USA
Weiss17 | Orosz I | University of California Los Angeles, USA
Weiss17 | Gibbs M | University of California Los Angeles, USA
Weiss17 | Schrader L | University of California Los Angeles, USA
Weiss17 | Lerner J | University of California Los Angeles, USA
Weiss17 | Cheng CK | University of California Los Angeles, USA
Weiss17 | Chang E | University of California Los Angeles, USA
Weiss17 | Rajaraman R | University of California Los Angeles, USA
Weiss17 | Keselman I | University of California Los Angeles, USA
Weiss17 | Churchman P | University of California Los Angeles, USA
Weiss17 | Bower-Baca C | University of California Los Angeles, USA
Weiss17 | Numis AL | University of California Los Angeles, USA
Weiss17 | Ho MG | University of California Los Angeles, USA
Weiss17 | Rao L | University of California Los Angeles, USA
Weiss17 | Bhat A | Thomas Jefferson University, USA
Weiss17 | Suski J | Thomas Jefferson University, USA
Weiss17 | Asadollahi M | Thomas Jefferson University, USA
Weiss17 | Ambrose T | Thomas Jefferson University, USA
Weiss17 | Fernandez A | Thomas Jefferson University, USA
Weiss17 | Nei M | Thomas Jefferson University, USA
Weiss17 | Skidmore C | Thomas Jefferson University, USA
Weiss17 | Mintzer S | Thomas Jefferson University, USA
Weiss17 | Eliashiv DS | Thomas Jefferson University, USA
Weiss17 | Mathern GW | University of California Los Angeles, USA
Weiss17 | Nuwer MR | University of California Los Angeles, USA
Weiss17 | Sperling M | Thomas Jefferson University, USA
Weiss17 | Engel Jr J | University of California Los Angeles, USA
Weiss17 | Stern JM | University of California Los Angeles, USA
Roy17 | Roy V | Jabalpur Engineering College, India
Roy17 | Shukla S | Jabalpur Engineering College, India
Roy17 | Shukla PK | University Institute of Technology, India
Roy17 | Rawat P | Truba Group of Institute, India
Melman16 | Melman T | Weill cornell Medical College, USA
Melman16 | Victor JD | Weill cornell Medical College, USA
Oliveira16 | Oliveira AS | University of Michigan, USA
Oliveira16 | Schlink BR | University of Michigan, USA
Oliveira16 | Hairston WD | Aberdeen Proving Ground, USA
Oliveira16 | König P | University of Osnabrück, Germany
Oliveira16 | König P | University Medical Center Hamburg-Eppendorf, Germany
Oliveira16 | Ferris DP | University of Michigan, USA
Cheveigne16 | Cheveigne A | CNRS, France
Cheveigne16 | Cheveigne A | Ecole Normale Superieure, France
Cheveigne16 | Cheveigne A | UCL Ear Institute, UK
Costa16 | Costa A | Miguel Hernandez University, Spain
Costa16 | Salazar-Varas R | Center of Research and Advanced Studies, Mexico
Costa16 | Ubeda A | Miguel Hernandez University, Spain
Costa16 | Azorin JM | Miguel Hernandez University, Spain
Nathan16 | Nathan K | University of Houston, USA
Nathan16 | Nathan K | The Houston Methodist Research Institute, USA
Nathan16 | Contreras-Vidal JL | University of Houston, USA
Nathan16 | Contreras-Vidal JL | The Houston Methodist Research Institute, USA
Zou16 | Zou Y | University of Texas at Dallas, USA
Zou16 | Nathan V | University of Texas at Dallas, USA
Zou16 | Jafari R | University of Texas at Dallas, USA
Gordon15 | Gordon SM | DCS Corporation Alexandria, USA
Gordon15 | Lawhern V | US Army Research Laboratory, USA
Gordon15 | Passaro AD | DCS Corporation Alexandria, USA
Gordon15 | McDowell K | US Army Research Laboratory, USA
Snyder15 | Snyder KL | University of Michigan, USA
Snyder15 | Snyder KL | University of Minnesota Duluth, USA
Snyder15 | Kline JE | University of Michigan, USA
Snyder15 | Huang HJ | University of Michigan, USA
Snyder15 | Ferris DP | University of Michigan, USA
Snyder15 | Ferris DP | University of Michigan, USA
Kline15 | Kline JE | University of Michigan, USA
Kline15 | Huang HJ | University of Michigan, USA
Kline15 | Snyder KL | University of Michigan, USA
Kline15 | Ferris DP | University of Michigan, USA
Kline15 | Ferris DP | University of Michigan, USA
Raduntz15 | Radüntz T | Federal Institute for Occupational Safety and Health, Germany
Raduntz15 | Scouten J | Federal Institute for Occupational Safety and Health, Germany
Raduntz15 | Hochmuth O | Humboldt-Universität zu Berlin, Germany
Raduntz15 | Meffert B | Humboldt-Universität zu Berlin, Germany
Mariani15 | Mariani S | Harvard University, USA
Mariani15 | Borges AFT | VU University Amsterdam, Netherlands
Mariani15 | Henriques T | Harvard University, USA
Mariani15 | Goldberger AL | Harvard University, USA
Mariani15 | Goldberger AL | Beth Israel Deaconess Medical Center, USA
Mariani15 | Costa MD | Harvard University, USA
Mariani15 | Costa MD | Beth Israel Deaconess Medical Center, USA
Zhang15 | Zhang C | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Tong L | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Zeng Y | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Jiang J | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Bu H | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Yan B | China National Digital Switching System Engineering and Technological Research Center, China
Zhang15 | Li J | China National Digital Switching System Engineering and Technological Research Center, China
Bulea14 | Bulea TC | National Institutes of Health, USA
Bulea14 | Bulea TC | University of Houston, USA
Bulea14 | Prasad S | University of Houston, USA
Bulea14 | Kilicarslan A | University of Houston, USA
Bulea14 | Contreras-Vidal JL | University of Houston, USA
Chen14 | Chen X | Hefei University of Technology, China
Chen14 | Chen X | University of British Columbia, Canada
Chen14 | Liu A | University of British Columbia, Canada
Chen14 | Peng H | Hefei University of Technology, China
Chen14 | Ward RK | University of British Columbia, Canada
Yilmaz14 | Yilmaz G | Koç University School of Medicine, Turkey
Yilmaz14 | Ungan P | Koç University School of Medicine, Turkey
Yilmaz14 | Sebik O | Koç University School of Medicine, Turkey
Yilmaz14 | Ugincius P | Lithuanian University of Health Sciences, Lithuania
Yilmaz14 | Türker KS | Koç University School of Medicine, Turkey
Reis14 | Reis PM | Friedrich-Alexander-University Erlangen-Nuremberg, Germany
Reis14 | Hebenstreit F | Friedrich-Alexander-University Erlangen-Nuremberg, Germany
Reis14 | Gabsteiger F | Friedrich-Alexander-University Erlangen-Nuremberg, Germany
Reis14 | von Tscharner V | University of Calgary, Canada
Reis14 | Lochmann M | Friedrich-Alexander-University Erlangen-Nuremberg, Germany
Boudet14 | Boudet S | University Catholic of Lille, France
Boudet14 | Boudet S | Unit´e de Traitement de Signaux Biomedicaux, France
Boudet14 | Peyrodie L | Unit´e de Traitement de Signaux Biomedicaux, France
Boudet14 | Peyrodie L | Hautes Etudes d’Ingenieur, France
Boudet14 | Szurhaj W | Clinical Neurophysiology service of CHR of Lille, France
Boudet14 | Bolo N | Harvard Medical School Boston, USA
Boudet14 | Pinti A | Universit´e d’Orleans, France
Boudet14 | Gallois P | University Catholic of Lille, France
Boudet14 | Gallois P | Unit´e de Traitement de Signaux Biomedicaux, France
Boudet14 | Gallois P | Groupe Hospitalier de l’Institut Catholique Lillois, France
Li13 | Li P | University of Electronic Science and Technology of China, China
Li13 | Xu P | University of Electronic Science and Technology of China, China
Li13 | Zhang R | University of Electronic Science and Technology of China, China
Li13 | Guo L | University of Electronic Science and Technology of China, China
Li13 | Yao D | University of Electronic Science and Technology of China, China
Lawhern13 | Lawhern V | University of Texas-San Antonio, USA
Lawhern13 | Hairston WD | US Army Research Laboratory, USA
Lawhern13 | Robbins K | University of Texas-San Antonio, USA
Nottage13 | Nottage JF | King's College, UK
Nottage13 | Morrison PD | King's College, UK
Nottage13 | ffytche DH | King's College, UK
Yong12 | Yong X | University of British Columbia, Canada
Yong12 | Fatourechi M | University of British Columbia, Canada
Yong12 | Ward RK | University of British Columbia, Canada
Yong12 | Birch GE | University of British Columbia, Canada
Yong12 | Birch GE | Neil Squire Society, Canada
Lau12 | Lau TM | University of Michigan, USA
Lau12 | Lau TM | US Army Research Laboratory, USA
Lau12 | Gwin JT | University of Michigan, USA
Lau12 | McDowell K | US Army Research Laboratory, USA
Lau12 | Ferris DP | University of Michigan, USA
Lawhern12 | Lawhern V | University of Texas-San Antonio, USA
Lawhern12 | Hairston WD | US Army Research Laboratory, USA
Lawhern12 | McDowell K | US Army Research Laboratory, USA
Lawhern12 | Westerfield M | University of California San Diego, USA
Lawhern12 | Robbins K | University of Texas-San Antonio, USA
Winkler11 | Winkler I | Technische Universität Berlin, Germany
Winkler11 | Haufe S | Technische Universität Berlin, Germany
Winkler11 | Tangermann M | Technische Universität Berlin, Germany
Gwin10 | Gwin JT | University of Michigan, USA
Gwin10 | Gramann K | University of Michigan, USA
Gwin10 | Makeig S | University of California San Diego, USA
Gwin10 | Ferris DP | University of Michigan, USA
McMenamin10 | McMenamin BW | Center for Cognitive Science, USA
McMenamin10 | Shackman AJ | University of Wisconsin, USA
McMenamin10 | Maxwell JS | US Army Research Laboratory, USA
McMenamin10 | Bachhuber DRW | University of Wisconsin, USA
McMenamin10 | Koppenhaver AM | University of Wisconsin, USA
McMenamin10 | Greischar LL | University of Wisconsin, USA
McMenamin10 | Davidson RJ | University of Wisconsin, USA
Mosher09 | Mosher JC | Cleveland Clinic Neurological Institute, USA
Mosher09 | Hämäläinen MS | Massachusetts General Hospital, USA
Mosher09 | Pantazis D | University of Southern California, USA
Mosher09 | Hui HB | University of Southern California, USA
Mosher09 | Burgess RC | Cleveland Clinic Neurological Institute, USA
Mosher09 | Leahy RM | University of Southern California, USA
McMenamin09 | McMenamin BW | University of Minnesota, USA
McMenamin09 | Shackman AJ | University of Wisconsin, USA
McMenamin09 | Maxwell JS | University of Wisconsin, USA
McMenamin09 | Greischar LL | University of Wisconsin, USA
McMenamin09 | Davidson RJ | University of Wisconsin, USA
Delorme04 | Delorme A | University of California San Diego, USA
Delorme04 | Makeig S | University of California San Diego, USA
Iriarte03 | Iriarte J | Private University of Navarra, Spain
Iriarte03 | Urrestarazu E | Private University of Navarra, Spain
Iriarte03 | Valencia M | Private University of Navarra, Spain
Iriarte03 | Valencia M | Public University of Navarra, Spain
Iriarte03 | Alegre M | Private University of Navarra, Spain
Iriarte03 | Malanda A | Public University of Navarra, Spain
Iriarte03 | Viteri C | Private University of Navarra, Spain
Iriarte03 | Artieda J | Private University of Navarra, Spain
Jung01 | Jung TP | University of California San Diego, USA
Jung01 | Jung TP | Salk Institute, USA
Jung01 | Makeig S | University of California San Diego, USA
Jung01 | Makeig S | Salk Institute, USA
Jung01 | McKeown MJ | Duke University, USA
Jung01 | Bell AJ | University of California San Diego, USA
Jung01 | Bell AJ | Salk Institute, USA
Jung01 | Lee TW | University of California San Diego, USA
Jung01 | Lee TW | Salk Institute, USA
Jung01 | Sejnowski TJ | University of California San Diego, USA
Jung01 | Sejnowski TJ | Salk Institute, USA

## papers_annotated
Reference | Index | Title | Authors | Journal/Book | reader | mental strategy | objective | data acquisition | sample (size, sex, age group, health status) | #sample | evaluation metrics | framework | open code (yes, no, -) | Innovation / Finding | further work suggested | Additional information
Hossain22 | 8 | Motion Artifacts Correction from Single-Channel EEG and fNIRS Signals Using Novel Wavelet Packet Decomposition in Combination with Canonical Correlation Analysis | Hossain MS, Chowdhury MEH, Reaz MBI, Ali SHM, Bakar AAA, Kiranyaz S, Khandakar A, Alhatou M, Habib R, Hossain MM. | Sensors (Basel) | mst | - | proposes 2 methods: (i) Wavelet packet decomposition (WPD) and (ii) WPD in combination with canonical correlation analysis (WPD-CCA), for motion artifact correction from single-channel EEG and compares it to multiple existing variants in benchmark tests | Single channel EEG with ground truth reference (second EEG + accelerometer used as ground truth). 23 sets of EEG recordings, sampled at 2048 Hz, later downsampled to 256 Hz | collected from six patients in four different sessions | 6 | SNR, average percentage reduction in artifacts | Matlab | no | Proposed WPD algorithms were investigated by 18 different approaches / modalities where four different wavelet packet families namely Daubechies, Symlet, Coiflet, and Fejer-Korovkin wavelet packet families | The paper describes the disadvantages of many existing methods like e.g. DWT, EMD, EEMD, VMD, DWT-ICA, EMD-ICA, EEMD-ICA, EMD-CCA, EEMD-CCA, VMD-CCA,.. BUT the paper does not show direct comparision of performance of these existing methods and the proposed variants of WPD and WPD-CCA | Would benefit from largely comparing DWT, EMD, EEMD, VMD, DWT-ICA, EMD-ICA, EEMD-ICA, EMD-CCA, EEMD-CCA, VMD-CCA with proposed WPD and WPD-CCA methods
Abu22 | 10 | Improved Cognitive Vigilance Assessment after Artifact Reduction with Wavelet Independent Component Analysis | Abu Farha N, Al-Shargie F, Tariq U, Al-Nashash H. | Sensors (Basel) | mst | attention strategy | Vigilance level assessment (alertness VS vigilance decrement) while under the effect of motion artifacts. Comparison of wICA and ICA as preprocessing for vigilance level classification | 64-Ag/AgCl scalp EEG gel-electrodes (ANT Neuro EEG system), at a sampling rate of 500 Hz; The AFz electrode was set as the system ground and electrodes M1 and M2 were used as reference electrodes | nine healthy volunteer students from the American University of Sharjah, aged 24.5 (+- 5.5 years). Defined criteria for participants: normal hearing, normal or corrected-to-normal vision, no history of psychiatric or cognitive disorders, no symptoms of drug addiction or abuse, and no intake of long-term medications | 9 | Accuracy, Specificity, Sensitivity | - | - | Wavelet ICA can lead to improved classification accuracy when applied to vigilance level assessment. The wICA approach reduced the distortion in the signal and was able to recover a substantial neural signal that was lost within artifacts. Delta, theta, and alpha bands showed high classification accuracy between the two defined mental states (alertness and vigilance decrement) | - | -
Flo22 | 16 | Automated Pipeline for Infants Continuous EEG (APICE): A flexible pipeline for developmental cognitive studies | Fló A, Gennari G, Benjamin L, Dehaene-Lambertz G. | Dev Cogn Neurosci | cs | attention strategy | (1) SSAEP in asleep neonates, (2) SSVEP in awake 5-month olds | 128-electrode net (Electrical Geodesics, Inc.) with 250 Hz in dataset (1) and 500 Hz in dataset (2) | healthy infants in two datasets: (1) 24 asleep neonates (11 males), (2) 26 awake 5-month olds (12 males), 80–140 trials per participant). | 24, 26 | proportion of retained epochs, standardized measurement error (SME) after Luck et al. (2021) | Matlab | yes | The APICE pipeline is proposed (Automatic Pipeline for Infants Continuous EEG) and validated in sevaral forms. | -
Kumaravel22 | 17 | NEAR: An artifact removal pipeline for human newborn EEG data | Kumaravel VP, Farella E, Parise E, Buiatti M. | Dev Cogn Neurosci | cs | SSVEP (frequency-tagging paradigm) | Analysis of EEG during visual stimuli | 125-electrode net (Electrical Geodesics, Inc.) with 250 Hz and impedance below 100kOhm | (1) 11 healthy neonates (6 males, 40 hours mean age), (2) 10 healthy neonates (6 males, 60 hours mean age) | 11, 10 | PDS, Frequency-Tagged Response, accuracy, F1-score | Matlab | yes | The proposed method is much better than MADE and HAPPE for infant EEG data | combining ICA with the ASR-based NEAR method
Soghoyan21 | 35 | A Toolbox and Crowdsourcing Platform for Automatic Labeling of Independent Components in Electroencephalography | Soghoyan G, Ledovsky A, Nekrashevich M, Martynova O, Polikanova I, Portnova G, Rebreikina A, Sysoeva O, Sharaev M. | Front Neuroinform | cs | attention strategy | Platform for the collection of EEG data and experts to label IC-components, Automated ML pipeline for the classification | 31-channel EEG NeuroTravel amplifier (EB Neuro, Italy), 10-10 system | 20 typical children with age between 5 and 14 years | 20 | AUC (of the ROC), Precision, F1-score | Python | yes | Low inter-expert consistency (like in the literature) | Re-evaluate the artifact classes, improve the interface and use more complex models like CNN | This improves the backdrawn of badly labelled EEG epochs, and could lead to a superior ML pipeline for automated IC-classification.
Karpiel21 | 40 | The Influence of Filters on EEG-ERP Testing: Analysis of Motor Cortex in Healthy Subjects | Karpiel I, Kurasz Z, Kurasz R, Duch K. | Sensors (Basel) | mst,cs | motor execution | participants were shown three numbers and they had to press with one of three respective fingers | 6 selected channels of 64-electrodes with ANT Neuro hardware and 10-20 system, sampled at 256Hz | 7 healthy right-handed subjects, aged between 20 and 35, 450 right-handed and 300 left-handed stimulus-response iterations per subject | 7 | latency, amplitude | Python | - | There is hardly any difference between the filters (FIR, IIR, FFT, NOTCH), the channels O1, O2 had the lowest ERP latency (<200ms) | There should be more work on comparing multiple pipelines and methods | No significant results, the channels O1, O2 had the lowest ERP latency (<200ms)
Beach21 | 47 | Motion artefact removal in electroencephalography and electrocardiography by using multichannel inertial measurement units and adaptive filtering | Beach C, Li M, Balaban E, Casson AJ. | Healthc Technol Lett | cs | motor execution | IMU on each active electrode. Using AF to reduce motion artifacts from head swinging directly in the electrode. | 4 active gel electrodes in the 10-20 system with 220Hz | 1 volunteer | 1 | SNR | Matlab | - | AF in the stage of improving the single-channel signal with electrode-specific accelerometer and gyroscope data improves the signal for lower SNR, but not for higher SNR. | More data has to be collected. Automatically detect the SNR to decide if to use AF and the acc or gyroscope signal. An online implementation is planned. Multi-channel AF should be investigated. | The focus is only AF in the active electrode. Not useful in the current form, but maybe they have improved the electrodes as ARM is a partner. They use a Butterworth filter and then integrate the signal, which is problematic (Bessel should be used)
Dai21 | 52 | A Fast Approach to Removing Muscle Artifacts for EEG with Signal Serialization Based Ensemble Empirical Mode Decomposition | Dai Y, Duan F, Feng F, Sun Z, Zhang Y, Caiafa CF, Marti-Puig P, Solé-Casals J. | Entropy (Basel) | cs | motor execution | improve the performance of EEMD while remaining same quality of the results for separating artifacual components for EEG data with added EMG artifacts | 22 EEG Ag/AgCl electrodes with additional 22 EMG electrodes and 3 EOG, 250 Hz samle rate, visual removement of artifactual epochs to get clean ground trues, for the EMG data, both sides of the forearm with Telemyo DTS, 250 Hz sample rate, channelwise weighted addition to get data with different SNR. Each subject has 288 trials with 6s each | 9 healthy subjects aged 22 to 25 for EMG, EEG from a preexisting dataset, semi-simulated data | 8 | RRMSE (relative RMSE: normalized with RMS(X-EEG), correlation coefficient, runtime | Matlab | - | sEEMD needs half the time of EEMD with comparable results, CCA performs better than ICA | - | a potential good way to preduce massive data for the pretraining of a very large NN is to add EMG data to visually cleaned EEG data
Tost21 | 54 | Choosing Strategies to Deal with Artifactual EEG Data in Children with Cognitive Impairment | Tost A, Migliorelli C, Bachiller A, Medina-Rivera I, Romero S, García-Cazorla Á, Mañanas MA. | Entropy (Basel) | cs | - | Children with Rett syndrome were recorded and their EEG epochs classified as artifact contaminated and grouped into symptomatic categories | 20 active electrodes (Starstim 20 wireless device, Neuroelectrics) in the 10-20 system, additional accelerometer and eye-tracking device, 84 records in total, 500Hz | 29 patients with Rett Syndrom (levels 1 to 4), 3 to 19 years, different cognitive and walking abilities | 29 | precision, recall, F1-score, accuracy | - | - | Parallel streams within the pipeline with different preprocessing with subsequent artifact classification increases the performance of the task | apply ICA after the removal of highly artifactual time-segments | The parallelization of the pipeline into multiple streams with different preprocessing could be interesting for us.
Kaur21 | 56 | EEG Artifact Removal System for Depression Using a Hybrid Denoising Approach | Kaur C, Singh P, Sahni S. | Basic Clin Neurosci | mst |  | proposes a novel denoising method based on Empirical Mode Decomposition (EMD) with Detrended Fluctuation Analysis (DFA) and wavelet packet transform. DFA is used as the mode selection criteria | sampling rate of data is 256 Hz, device is not mentioned directly in paper | 34 MDD (Major Depressive Disorder) patients and 30 healthy subjects. Open dataset reused - provided by Hospital Universiti Sains Malaysia (HUSM) | 34 | SNR, MAE | - | - | Proposes EMD-DFA-WPD pipeline and compares SNR and MAE of its results to 3 existing pipelines DWT, EMD and EMD-DWT. Additionally, uses SVM and RF for classification of depression states (not entered as pipelines here) | In the future, more efficient algorithms can be developed by increasing the levels of decompositions and considering other partially variational algorithms to decrease the problem of mode mixing by EMD | -
Rosanne21 | 73 | Adaptive Filtering for Improved EEG-Based Mental Workload Assessment of Ambulant Users | Rosanne O, Albuquerque I, Cassani R, Gagnon JF, Tremblay S, Falk TH. | Front Neurosci | mst | motor execution | improve mental workload assessment while on a treadmill or stationary bike. Total low physical activity and high pyhisical activity setup for a 2 class mental workload classification problem | Neurolectrics Enobio 8-channel portable headset with positions Fp1, Fp2, AF7, AF8, T9, T10, P3, P4; sampling rate 500 Hz and later downsampled to 250 Hz | 48 participants (23 females, 25 males); average age was 27.4 ± 6.6 years old; of those participants 22 utilized a treadmill during the experiment and 26 a stationary bike | 48 | accuracy (for classification) | MATB-II: Microsoft Visual Studio.NET (VS.NET) | - | use of an adaptive filtering scheme with methods ASR, ADJUST, HAPPE, ASR_ADJUST, ASR_wICA and wICA can be used to remove movement artifacts from EEG signals and increase robustness of simultaneous mental workload assessment | - | best combined with HAPPE and can result in 97% mental workload prediction accuracy for both low and high physical activity conditions
Nahmias21 | 82 | Quantifying Signal Quality From Unimodal and Multimodal Sources: Application to EEG With Ocular and Motion Artifacts | Nahmias DO, Kontson KL. | Front Neurosci | mst | Motor imagery, Mental arithmetic | novel methods for unimodal recording methods (here e.g. ocular artifacts) and multimodal recording methods (here e.g. motion artifacts). In case of motion artifacts a deep learning CNN is presented | recorded with 58 electrode array (10-20 system) while EOG were recorded with four electrodes placed above and below each eye. Additionally motion data was captured simulataneously with IMU | EEG and EOG recordings (sampled at 100 Hz) as well as motion capture from eleven subjects walking on a treadmill for 6 min (pre-existing dataset from University of Houston). Subjects walked at one, two, three, and four miles-per-hour. | 11 | accuracy (for classification) | Matlab | yes | The findings of the unimodal recording method of this paper are not interesting for this list, but the multimodal method results are. The multimodal approach can be applied to data with unimodal source anyways given enough data for learning and appropiate model design. Given enough data for learning the DCNN reaches good classification accuracies. | More deep learning approaches should be investigated for EEG artifact removal. This is one of the few papers that proposes a DCNN for motion artifact removal. | -
Saba-Sadiya21 | 86 | Unsupervised EEG Artifact Detection and Correction | Saba-Sadiya S, Chantland E, Alhanai T, Liu T, Ghassemi MM. | Front Digit Health | cs | attention strategy | SSVEP: (1) orientation task, (2) color task | 32-electrode actiCHamp cap at 1000 Hz. Data from Saba-Sadiya20 | in both datasets 7 subjects (4 males) with a total of ~10,000 EEG trials. All subjects reported a normal or corrected to normal vision | 7 | f1-score, Cohen's Kappa, Accuracy | Matlab, Python | yes | A novel EEG noise-reduction method based on representation learning. Artifact detection based on 58 extracted features and statistical or representation learning methods. | Validate the method on additional tasks and datasets. More complex elements for the network Architecture
Li21 | 89 | Electromyogram (EMG) Removal by Adding Sources of EMG (ERASE)-A Novel ICA-Based Algorithm for Removing Myoelectric Artifacts From EEG | Li Y, Wang PT, Vaidya MP, Flint RD, Liu CY, Slutzky MW, Do AH. | Front Neurosci | mst | motor execution | Proposed a new method (ERASE) that relies on recorded or simulated EMG data input for correcting/removing contamination from EEG data | 64-channel EEG cap (ActiCap, Brain Products, Gilching, Germany) acquired by two linked NeXus-32 systems (Mind Media,Herten,Netherlands); EMG was recorded using a MP150 system (BIOPAC,Goleta CA) | Simulated data tests + real data collection from 8 healthy subjects; all subjects were male and between 20 and 50 years old; a total of 12 sessions (120 trials) were performed across all subjects | 8 | z-scored power | - | - | New method, were additional channels of real EMG from neck and head muscles (reference artifacts) were added as inputs to ICA in order to “force” the most power from EMG artifacts into a few independent components (ICs). The ICs containing EMG artifacts were identified and rejected automatically | Results suggest that using reference EMG artifacts can force ICA to learn and detect the EMG artifacts by forcing the contaminant EMG within EEG into a minimal number of ICs. Comparison of multiple EMG-added methods to test this assumptions of the authos is needed. | Has been validated with simulated data primarily; only study done is with small sample size (8). Needs more evaluation with real data coming from a more diverse group of participants
Anders20 | 106 | The influence of motor tasks and cut-off parameter selection on artifact subspace reconstruction in EEG recordings | Anders P, Müller H, Skjæret-Maroni N, Vereijken B, Baumeister J. | Med Biol Eng Comput | cs | motor execution | Optimization of ASR cut-off parameters on three different tasks (1) n-back mental task, (2) single-leg stance and (3) forward and backward walking | 64-channel EEG with Ag/AgCl active wet electrodes on an elastic cap (Easycap, Germany) in an extended 10-20 system and a wireless amplifier (Live Amp, Brain Products, Germany). Sampled at 500Hz and then downsampled to 250Hz | 5 healty females with mean age of 23.2 years. | 5 | Ratio of removed and reconstructed data, RELICA's quality and dipolarity index, reproducability using bootstrapping | Matlab | - | ASR should be used with a cut-off frequency of at least 10 std, better 20 for tasks with many movement. | Investigate the effect of movement task on the optimal cut-off parameter more
Zhang20 | 117 | Fully organic compliant dry electrodes self-adhesive to skin for long-term motion-robust epidermal biopotential monitoring | Zhang L, Kumar KS, He H, Cai CJ, He X, Gao H, Yue S, Li C, Seet RC, Ren H, Ouyang J. | Nat Commun | mst,cs | motor execution | open and close eyes; finger motion EMG | new PEDOT:PSS electrodes VS established Ag/AgCl electrodes | single patient with atrial fibrillation, rest unknown | 1 | - | Matlab | - | new conductive polymer dry electrode (PEDOT:PSS) with lowskin-contact impedance and noise, good stretchability and conductivity | - | is only tested with small sample size and probably needs more testing and direct gold standard comparison studies. The results do not show how this approach can mitigate motion artifacts, but they indicate that the electrodes could be used longer and even a week with a smaller increase in noise compared to standard gel electrodes.
Leach20 | 126 | Adjusting ADJUST: Optimizing the ADJUST algorithm for pediatric data using geodesic nets | Leach SC, Morales S, Bowers ME, Buzzell GA, Debnath R, Beall D, Fox NA. | Psychophysiology | cs | - | They optimize the adjusted ADJUST algorithm for features classification to children and infancts | 64-channel HydroCel Geodesic Sensor net and EGI Netstation software, sampled at 500Hz, amplified with EGI NetAmps 300 For the testing set, a 128-electrode Net was subsetted to 64-channels to be consistent. | Training data: 96 children (age 3 to 9 years). Test data: (1) 10 children (4 males) aged 7 to 9, (2) 10 adults (3 males) aged 20 to 22 years. (3) 10 infants (5 males) aged 4 to 6 months | 96, 10, 10, 10 | Percentage of agreement, trials retained, reliability | Matlab, EEGLAB | yes | adjusted-ADJUST and ICLabel are both well suited for adults, but adjusted-ADJUST is better for children and infants | Train for more data, only data for 30 participants were labelled
Saini20 | 139 | Effective automated method for detection and suppression of muscle artefacts from single-channel EEG signal | Saini M, Satija U, Upadhayay MD. | Healthc Technol Lett | mst | - | Proposed a new and simple 3-step method for detection and suppression of muscle artifacts from single EEG channels | reused online publically available databases for evaluation: [1] Mendeley database, [2] epileptic Bonn database, [3] EEG during mental arithmetic tasks (EEGMAT), [4] examples of electromyograms, [5] cerebral vasoregulation in elderly with stroke (CVES) database | [1] not specified in paper; [2] five healthy subjects; [3] 36 subjects; [4] unknown number of healthy subjects; [5] not specified | 5 | percentage root mean square difference (PRD), RMSE | - | no | A new single EEG channel detection/reduction method is proposed that works using 3 steps: [1] variational mode decomposition (VMD) to get two modes, [2] rejection of MA components in the second mode based on zero crossings count thresholding, [3] Finally, in order to get a artifact-free EEG signal, the first mode is retained only if the presence of muscle artifacts is detected in the second mode; otherwise, the original signal is considered as artifact-free signal | Although the authors evaluate the proposed method against existing methods, an independant method comparison by another institute should be considered, as the overall superior performance of this method (additionally to it seemingly not having any real downsides) seems unlikely | High value paper, if proposed results are actually unbiased and not handpicked to show the strenghts of the new method. In the introduction section the authors analyze the weaknesses of existing methods, which is an interesting read in general
Mucarquer20 | 144 | Improving EEG Muscle Artifact Removal With an EMG Array | Mucarquer JA, Prado P, Escobar MJ, El-Deredy W, Zañartu M. | IEEE Trans Instrum Meas | mst | attention strategy | adding an EMG array to EEMD-CCA | clean data set: 32-channel of EEG recorded from two subjects involved in a passive listening task (auditory oddball paradigm during hypnosis), obtained from BNCI Horizon 2020 project Contamination data: EMG signals were recorded from a subject in resting state and in an exercise of vocalization, uttering a set of long syllables. | clean data set: two subjects contamination: one subject | 3 | RRMSE | Matlab | - | Independently of the number of muscular sources, a substantial performance improvement is expected by placing additional EMG electrodes over the face and neck muscles. The number of optimal EMG channels must be associated with the number of SMU units interfering the signal | Effects of EMG addition should be tested with other algortihms of the pipeline as well | -
Nordin20 | 145 | Faster Gait Speeds Reduce Alpha and Beta EEG Spectral Power From Human Sensorimotor Cortex | Nordin AD, Hairston WD, Ferris DP. | IEEE Trans Biomed Eng | mst | motor execution | Determine if walking speed affects human sensorimotor electrocortical dynamics using novel mobile high-density EEG (simultaneously recorded human electrocortical signals and isolated motion artifacts using pairs of mechanically coupled and electrically independent electrodes). Compared artifact removal methods during treadmill walking at four speeds (0.5, 1.0, 1.5, and 2.0 m/s). | 128-channel gel-based EEG cap with BioSemi ActiveTwo sensors (40 mechanically coupled and inverted noise-only electrodes) and the location of each scalp electrode was measured with a Zebris digitizing system. Additionally used 8-neck EMG electrodes (BioSemi ActiveTwo) | nine healthy subjects (6 male, 3 female, mean age 27 ± 4 years). It is important to note that in the result section of evaluating the pipeline algortihms only 5-6 subjects data have been used | 9 | Relative power (db), scalp map, dipole locations and ERSP plots | MATLAB, EEGLAB, ... | - | Left and right sensorimotor alpha/beta spectral power increased while contralateral limb single support and push off, and decreased during contralateral limb swing at each speed. At faster walking speeds, sensorimotor spectral power fluctuations were less pronounced across the gait cycle with reduced alpha and beta power (p<0.05) compared to slower speeds. Isolated noise recordings and neck EMG spectral power fluctuations matched gait events and showed broadband spectral power increases at faster speeds. | Proposed dual layer system should be further investigated in upcoming new studies. Findings of walking could be extended for running, sports and task execution scenarios. | Very high value findings (!!!) for understanding walking at various speeds and how different phases of walking impact spectral power
Liu19 | 156 | Removal of EMG Artifacts from Multichannel EEG Signals Using Combined Singular Spectrum Analysis and Canonical Correlation Analysis | Liu Q, Liu A, Zhang X, Chen X, Qian R, Chen X. | J Healthc Eng | cs | - | Compare SSH-CCA with other well successful pipelines in order to remove EMG parts from a semisimulated data of clean EEG and EMG. | (1) 19-channel EEG at 500Hz + lambda* EMG data at 500Hz. Added with different fraction to achieve different SNR, (2) real EEG with 21 channels at 250Hz | (1: EEG) 20 healthy subjects, (2: EMG) healty subjects, (3: real EEG) 1 subject | 20, 1 | RRMSE (relative RMSE), accuracy, PDS, visual, time | Matlab | - | The combination of single channel decomposition with BSS seems to be good. SSA-CCA is superior to other combinations of EEMD, SSA, ICA and CCA. Applying SSA twice is enough decomposion (for theEMG part) | They might improve their proposed SSA-CCA to satisfy the needs of removing muscle artifacts from few-channel EEG signals | Very interesting introduction, they criticize ICA for muscle artifacts a lot. The method to generate semisimulated data based on clean EEG and add EMG to it seems to be very nice for us, as we get ground truth. It might not be good to use a mixing matrix that sometimes give no artifacts at all to the EEG, as it eases the learning of the method to much.
Benda19 | 162 | Peak Detection with Online Electroencephalography (EEG) Artifact Removal for Brain-Computer Interface (BCI) Purposes | Benda M, Volosyak I. | Brain Sci | mst | motor imagery, motor execution | novel automatic online artefact removal technique using peak alpha frequency detection; filtered and unfiltered comparison by using an existing artefact removal method (multi-channel Wiener Filter). Used to improve classification of user fatigue | 14 Ag/AgCl gel electrodes, recorded at sampling rate 256 Hz. Used biosignal amplifier was g.USBamp (Guger Technologies, Graz, Austria); | Fifteen healthy users (eight female) participated. All of them were students of the Rhine-Waal University of Applied Sciences. The average age (SD) was 23.8 (2.9) years. Three phases of data acquisition using motor imagery and motor execution of eyes and head movements (10 tasks total) | 15 | peak detection efficacy, peak detection accuracy | - | - | Beneficial for peak alpha f detection in cases of eye movement and head movement. Results were not compared to other existing methods | Presented novel method needs a method comparison study | -
Val-Calvo19 | 163 | Optimization of Real-Time EEG Artifact Removal and Emotion Estimation for Human-Robot Interaction Applications | Val-Calvo M, Álvarez-Sánchez JR, Ferrández-Vicente JM, Fernández E. | Front Comput Neurosci | cs | attention strategy | Subjects were employed a set of 15 emotionally tagged videos of three categories. A pipeline was optimized for emotion estimation with a focus on the classifacator and real time ability | Referenced the SEED database, 64 channels in the 10-20 system. | SEED database with 15 subjects and 3 sessions of 15 videos of approximately 180s each. | 15 | RMSE, correlation coefficient, mutual information, runtime, accuracy | Python | - | Feature smoothing is an important part for preprocessing and especially Savitzky-Golay is better suited than linear dynamic system. Focus on less electrodes (8) and features (50-100) is a good tradeoff between runtime and performance. Validation scheme has to be noted, they differ inter-trial, inter-session and inter-subject validation | Complex deep learning strategies must be implemented. Further research in HRI (human-robot interaction). | The setup of the experiments was very neat, but they should have been providing more information on their tests.
Saavedra19 | 171 | Wavelet-Based Semblance Methods to Enhance the Single-Trial Detection of Event-Related Potentials for a BCI Spelling System | Saavedra C, Salas R, Bougrain L. | Comput Intell Neurosci | mst | attention strategy, P300 | Two new methods are developed for single trial ERP detection: “multichannel EEG thresholding by similarity” (METS) + “semblance-based ERP window selection” (SEWS). | 10 electrodes were recorded (Fz, C3, Cz, C4, P3, Pz, P4, PO7, PO8, and Oz). Signals were recorded at 256 Hz using a g.tec g.USBamp EEG amplifier, a right ear reference, and a right mastoid ground. | P300 dataset recorded by the Neuroimaging Laboratory of Universidad Autonoma Metropolitana (UAM), Mexico, using the P300 speller included in the BCI2000 platform. Contains 22 first-time healthy users, and all subjects have similar characteristics, such as sleep duration and age. | 22 | accuracy (for classification of letters in P300 speller) | - | - | Combining both algorithms (METS + SEWS) for denoising signals and selecting the time window has been compared to xDAWN, which is an efficient algorithm to enhance ERPs.This new wavelet-based semblance method performs better than xDAWN in the presence of artifacts | It is unclear if this new combination is working well for muscle artifacts in particular as the entire paper does never mention clearly the type of artifacts mitigated - needs evaluation paper for muscle artifacts | -
Mur19 | 194 | An Unsupervised Method for Artefact Removal in EEG Signals | Mur A, Dormido R, Duro N. | Sensors (Basel) | cs | - | The subjects were measured without any task | 64-channel Biosemi Active Two System, sampled at 256Hz, different types of artefacts were recorded | One subject and a duration of 26 seconds | 1 | visual | - | - | The proposed UAR pipeline can correct real and synthetic artifacts very well, but it depends on the scarcity of them. | - | It is very interesting, that their proposed UAR method works incredibly well on simulated artifacts. That could suggest, that muscle artifacts coincidences with multiple brain activities. The descriptions are not clear and some interpretations not explained
Blum19 | 196 | A Riemannian Modification of Artifact Subspace Reconstruction for EEG Artifact Handling | Blum S, Jacobsen NSJ, Bleichner MG, Debener S. | Front Hum Neurosci | cs | attention strategy | Subjects were standing or walking and data was acquired on a smartphone. They compare ASR with an adapted version to detect eye blinks with low latency. | 24-channel EEG measured and processed on a smartphone | 27 participants (7 male, mean age 23 years) | 27 | sensitivity, specificity, latency | Matlab, EEGLAB | yes | Measurement and processing on a smartphone. ASR with Riemann geometry is faster with similar results. | Evaluate under more complex artifacts. Incorporate IMU data | Very cool data acquition on a smartphone, very good for BCI-systems. They highlight the importance of subject-specific calibration data in rest situation.
Nordin19 | 207 | Human electrocortical dynamics while stepping over obstacles | Nordin AD, Hairston WD, Ferris DP. | Sci Rep | mst | attention strategy, motor execution | studying brain dynamics during visually guided locomotion (unexpected obstacles appearing on a treadmill) by using new mothod of artifact removal and novel dual-layer system | collected 128-EEG and 40-noise electrodes + 8 electromyography neck electrodes (176 total channels, 88-single electrode EEG and 40-dual electrode EEG channels), IMU data from forehead (128 Hz sample rate, APDM, Inc., Portland, OR), 10-camera system capturing lower extremity kinematics (100 Hz sample rate, Vicon, UK), and a force-instrumented treadmill (1000 Hz sample rate, Bertec, Columbus, OH) | Nine (6 male, 3 female) healthy, right limb dominant, subjects participated | 9 | spectral perturbation | Matlab, EEGLab | yes | By using combined hardware and signal processing solutions for motion artifact removal, the paper found it is possible to identify human brain activity while humans stepped over obstacles while walking and running | The authors concluded: "According to PubMed, there were over 2,800 studies on human EEG published in 2017, yet less than 1% were on mobile subjects." | Very high value paper for real life motion analysis, claiming it possible to use EEG for running and walking tasks. Could be a foundation for our work.
Grosselin19 | 216 | Quality Assessment of Single-Channel EEG for Wearable Devices | Grosselin F, Navarro-Sune X, Vozzi A, Pandremmenou K, De Vico Fallani F, Attal Y, Chavez M. | Sensors (Basel) | mst | motor exectution, motor imagery | presented new classification-based approach to effectively quantify artefact contamination and evaluated it on 3 labeled databases (2 real DBs + 1 with artificially generated artifacts) and 2 unlabeled databases. Classifier performance compared of LDA, Linear SVM, Euclidean and Weighted kNNs. Classifiers classify into HIGH-Q, MED-Q and LOW-Q (signal quality) | artBA: Acticap BrainProducts (GmbH, Gilching, Germany) system using 32 wet electrodes artMM: Melomind (myBrain Technologies, Paris, France) with two dry sensors on P3 and P4 positions publicDB: g.tec GAMMAsys system using 30 wet active electrodes (g.LADYbird) and two g.USBamp biosignal amplifiers (Guger Technolgies, Graz, Austria) 2 unlabeled DBs: Acticap BrainProducts and Melomind, myBrain Technologies | artBA: 3 subjects artMM: 21 subjects publicDB: 9 subjects 2 unlabeled DBs: 10 subjects | 3,21,9,10 | Accuracy, AUC | Matlab | - | results suggest that the presented approach is a good EEG quality checker in offline environments with either dry or wet EEG electrodes. It was noted by the authors in the conclusion section that optimization of subject-driven classifiers for longitudinal recordings (weeks, months) might increase classification performance. Results indicate that the proposed method is suitable for real-time applications dealing with embedded EEG in mobile environments, such as the monitoring of cognitive or emotional states, ambulatory healthcare systems or sleep stage scoring | Needs longer lasting studies for evaluation of longitudinal recordings performance increases | For internal ranking only databases artBA and artMM are evaluated as different classifiers are only compared there
Li18 | 228 | An Approach for Brain-Controlled Prostheses Based on a Facial Expression Paradigm | Li R, Zhang X, Lu Z, Liu C, Li H, Sheng W, Odekhe R. | Front Neurosci | cs | attention strategy | The subjects used a FE-BCI (facial expression) to send four types of commands to a robot arm. | Neuracle 8-channel wireless EEG at 1kHz, but resampling to 250 Hz, 10-20 system, impedances below 5kOhm | 18 healthy subjects (15 males), between 22 and 30 years | 18 | accuracy | - | - | Only four electrodes suffice in getting 81% accuracy in a four class problem. Less electrodes or more wavelet features reduce it. Sample Entropy and MEMD method was successful | More representative subjects (number, gender, health status), idle state for BCI control | He overcame the limited IC component problem by using MEMD. Used db-3 Wavelet transform to produce features. Might only works because the facial muscles induces characteristic EEG signal and they don't origin from the cortex.
Chu18 | 240 | A Decoding Scheme for Incomplete Motor Imagery EEG With Deep Belief Network | Chu Y, Zhao X, Zou Y, Xu W, Han J, Zhao Y. | Front Neurosci | mst | motor imagery | Improvement of recognition accuracy and stability for incomplete EEG signals when performing different motor imagery tasks (imagining left hand, right hand and foot movements) | Recorded data with sampling rate of 1 kHz using OmmiPlex Neural Data Acquisition System cap (Plexon Inc., USA) with 64 Ag/AgCI passive gel electrodes (Plexon Inc., USA). Electrodes with roughly 3 cm separation distance closely arranged according to international 10–20 system. | 9 healthy right-handed volunteers (all male, mean age 26.5 years, range from 25 to 28 years) with thin hair participated in the motor imagery experiments. All subjects without any history of neurological, psychiatric or cognitive disorders. No participant had any prior experience with motor imagery BCI experiments. | 9 | accuracy | - | - | Proposes a pipeline where they compare Lomb-Scargle Periodogram (LSP), FFT and Welch for feature extraction. Additionally compares learning strategies SVM and restricted Boltzmann machine (RBM) with each other. | According to authors: "comparative research should be conducted by applying the proposed method to three different segmentation strategies: (1) no segmentation, by applying the proposed method directly to the whole EEG trial, (2) uniform segmentation without overlapping, and 3) segmentation with different overlapping (sliding window)" | Comes with many results of direct comparison in feature extraction changes and classifier changes within the proposed pipeline
San-Martin18 | 241 | A Method for Simultaneous Evaluation of Muscular and Neural Prepulse Inhibition | San-Martin R, Zimiani MI, Noya C, Ávila MAV, Shuhama R, Del-Ben CM, Menezes PR, Fraga FJ, Salum C. | Front Neurosci | cs | - | The subjects were stimulated by two subsequent auditive stimuli that lead to sensory overload. | EEG and EMG were recorded simultaneously at 512Hz, dry-active electrode cap (BrainVision) with 11 scalp electrodes, 10-20 system. | 22 healthy subjects (15 males) without neurologic or psychatric disorder, mean age of 26 | 22 | - | Matlab | - | The first method to evaluate muscular and neural PPI (prepulse inhibition) simultaneously that is a problem with people with psychiatric disorder. More artefacts in the frontal electrodes | Larger PPI database, alternative for ICA, tests with schizophrenia patients | Not our topic but interesting how sensory overload affects some people more.
Sebek18 | 248 | Suppression of overlearning in independent component analysis used for removal of muscular artifacts from electroencephalographic records | Sebek J, Bortel R, Sovka P. | PLoS One | mst | - | Supression of effects of overlearning when using ICA on short signals to remove muscle artifacts. Supression is achieved by projecting an EEG record into several subspaces | The data from 111 electrode system with a reference placed on a forehead were digitized with 16 bit resolution and sampled at the frequency of 1024Hz. | 20 healthy subjects (10 male, 10 female) with age ranging from 19 to 37 years (25.4 ± 5.1 years). During the acquisition of EEG data, the participants were seated with their head supported by a headrest, and instructed to relax. | 20 | correlation coefficient, SNR | - | - | Work presents an ICA added method for muscular artifact removal that is not prone to overlearning when used on short signals. Therefore it is better suited to remove artifacts that are distinctly non-stationary and appear, disappear or change their origin within a time frame of a few seconds or less. The algorithm can remove artifacts only in the short segments where it is applied. If the artifacts are not detected, the original EEG is not distorted. If the composition of artifacts is changed, the algorithm can react quickly. | Authors noted that suggested methodology may work better with some other ICA algorithms than FastICA, but testing the performance with other ICA algorithms was out of the scope of this paper | Uses Moore-Penrose Pseudoinverse to solve the ICA component limitation problem: Matrix division requires quadratic matrices on mixing process (matrix dimensions being channels times ICs). Matrix division means creating the inverse of the matrix and perform a multiplication. Creating an inverse of a matrix is only possible for quadratic matrices. Therefore possible maximum ICs obtained by ICA methods is limited to the number of maximum channels. Channel size restricts max. components.
Arad18 | 259 | Performance-based approach for movement artifact removal from electroencephalographic data recorded during locomotion | Arad E, Bartsch RP, Kantelhardt JW, Plotnik M. | PLoS One | mst | motor execution | Presents method for removing motion artifacts from EEG signals recorded during treadmill (TM) and over-ground (OG) walking, as well as quantify the prevalence of MA in different locomotion settings | used 32-electrode wireless EEG system (eego sports 32 pro by ANT-Neuro, NL), which utilizes passive-wet electrodes, a 3-axis accelerometer (eego sports 32 pro by ANT-Neuro, NL) and bi-lateral force plates (R-Mill, ForceLink, NL) | study included 5 young, healthy adult (mean age ± SD: 30.5 ± 5.31) participants. | 5 | group average, SD | Matlab | - | Work observed a signal increase (e.g., neck EMG activations during heel-strikes) when walking speed increased, that was not caused by physical motion. MAs were identified and quantified in each component using a novel method that utilizes the participant's stepping frequency, derived from a forehead-mounted accelerometer. used benchmarking the EEG data by applying newly established metrics to quantify the success of the method. | - | Compared Preprocessed VS AMICA VS ASR at differnt speeds
Symeonidou18 | 265 | Effects of Cable Sway, Electrode Surface Area, and Electrode Mass on Electroencephalography Signal Quality during Motion | Symeonidou ER, Nordin AD, Hairston WD, Ferris DP. | Sensors (Basel) | mst | - | Evaluation on contribution to observed motion artifacts of mass and surface area of different commercially available electrodes. Used gelatin head phantom with embedded antennas as a ground truth signal | BioSemi Active Two system and active, wet, flat type Ag/AgCl electrodes (BSM, BioSemi, Amsterdam, The Netherlands). EEG was recorded at 512 Hz and a bandwidth between 0.01 and 256 Hz. | A robotic platform moved the phantom head through sinusoidal displacements at different frequencies (0–2 Hz). Paper also adresses and reuses multiple studies with participants from other researchers. |  | SNR | Matlab, EEGLab | - | Main findings: A larger electrode surface area can have a small but significant effect on improving EEG signal quality during motion and that cable sway is a major contributor to motion artifacts. Cable sway largely contributes to observed motion artifacts and resulted in a 10% overall decrease in the SNR. Wireless EEG system resulted in the lowest SNR decrease with increasing motion frequency and displacement compared to the wired systems. Pre-amplification of the signal still results in significant signal quality decrease with cable motion. | Additional studies required on wireless systems examining signal quality during head motion in order to determine if wireless systems offer advantages over wired systems in terms of signal quality. In the future, wireless EEG systems might be able to substitute wired systems for locomotion experiments. Paper shows that there are still considerations concerning signal quality, user comfort, and electrode density that need to be addressed in future studies. | Makes good arguments in favour of wireless systems over wired systems. Setup was simulated with robot motion though, could benefit from real study with real participants
Stone18 | 268 | Automatic Removal of Physiological Artifacts in EEG: The Optimized Fingerprint Method for Sports Science Applications | Stone DB, Tamburro G, Fiedler P, Haueisen J, Comani S. | Front Hum Neurosci | mst | motor execution | Optimization and parameter tuning to enhance the previously introduced Fingerprint Method (Fingerprint Method VS Optimized Fingerprint Method). Used a sports performance paradigm where participants performed an endurance cycling task. | Two systems used: [1] standard gel-based 128 channel electrode cap (Ag/AgCl electrodes) arranged in quasi-equidistant montage (Waveguard, Advanced Neuro Technologies., NL) [2] novel set of 97 dry multipin polyurethane electrodes with a Ag/AgCl coating arranged in equidistant montage (Fiedler 2015) | Two studies: [1] 17 male non-clinical volunteers (age 28 ± 4 years) participated in the acquisition of the artifact-rich training and testing EEG data that were used to optimize the Fingerprint Method. [2] 22 male athletes (age 27 ± 7 years) volunteered in the cycling performance experiment that provided the EEG datasets that were used to validate the Optimized Fingerprint Method. | 17,22 | accuracy (for classification) | Matlab, EEGLab | - | Unique sets of features are suitable for the distinct detection of artifact types and that the Optimized Fingerprint Method was able to correctly identify over 90% of the artifactual components with physiological origin present in the experimental data. | The sport application domain rarely is explored in EEG reasearch yet, although it is an ideal domain to investigate EEG under high muscle artifact contamination and in settings with sweat noise | Artifacts categorized and presented by types: (1) Eye Blink, (2) Eye Movement, (3) Myogenic Only Myogenic artifact results are entered and displyed in this list, eye related results were ignored
Frolich18 | 272 | Removal of muscular artifacts in EEG signals: a comparison of linear decomposition methods | Frølich L, Dowding I. | Brain Inform | mst,cs | motor execution | Comparision of 5 BSS methods during motor execution | 64 equidistant Ag/AgCl electrodes measured at 1000 Hz and downsampled to 200 Hz, overly noicy electrodes were rejected | 18 subjects in a simulated driving experiment, self-paced right foot movement task with rythmic head movement as noice source | 18 | ERD peak score (event-related desynchronisation, peak of the grand-average in the beta band), dipolarity (percentage of compoonents whose scalp maps can be explained as dipole in Delorme 2012) | EEGlab, Matlab | - | A high-pass filter with 14 Hz cuttoff increase the results for most ICA-algorithms, InfomaxICA yields the best results.
Gabard-Durnam18 | 273 | The Harvard Automated Processing Pipeline for Electroencephalography (HAPPE): Standardized Processing Software for Developmental and High-Artifact Data | Gabard-Durnam LJ, Mendez Leal AS, Wilkinson CL, Levin AR. | Front Neurosci | mst | - | Introduction of automated pipeline "Harvard Automated Processing Pipeline for EEG" (HAPPE). Describes each HAPPE processing step, performs an example analysis and shows that it outperforms 7 alternative approaches. Provides reference distribution of HAPPE’s data quality metrics in an 867 file dataset. | HAPPE supports 64 and 128 channel devices. Dataset of result section was recorded with HydroCel 128-channel Geodesic Sensor Nets (EGI, Eugene, OR), but only 39 channels were used for analyisis. There are many different other example recordings mentioned in the paper. | Used dataset includes EEGs from infants spanning 3 months to 36 months of age, and contains three participant groups as part of a larger, longitudinal study on emergence of Autism Spectrum Disorder. |  | average percent of artifact level retained, average percent of EEG signal retained | EEGlab, Matlab | yes | HAPPE removes more artifacts than all presented 7 alternative approaches (ICA, Manual, ASR, ADJUST, SASICA, FASTER, FASTER-MARA) while simultaneously preserving greater or equivalent amounts of EEG signal in almost all instances looked at in this paper. | Limitations of HAPPE should be considered for further research (Users must select a subset of channels for HAPPE processing as high-density EEG recordings, but dev samples most likely will not meet the data-length requirements for robust ICA decomposition). It is suggested to try to use HAPPE with PCA or with previous dimensionality reduction steps. | First occuronce and introduction of HAPPE; HAPPE software is freely available under the terms of the GNU General Public License at https://github.com/lcnhappe/happe.
Tamburro18 | 275 | A new ICA-based fingerprint method for the automatic removal of physiological artifacts from EEG recordings | Tamburro G, Fiedler P, Stone D, Haueisen J, Comani S. | PeerJ | mst,cs | motor execution | Feature-based classifcation of artifact types (eye blinks, eye movement, myogenic, cardiac) within IC | RefaExt (Advanced Neuro Technologies) sampled at 1024Hz , 128 Ag/AgCl electrodes both wet cap (Waveguard, ANT) with 128 gel electrodes (ECI-Electrogel, Electrocap International Inc.) and 97 dry Multipin Polyurethane electrodes (both ring-shaped) | For (1) eyeblinks twelve subjects, (2) eye movement 19 subjects, (3) myogenic 19 subjects that perform for example jaw movements. Only men with mean age mid twenties. All with gel and dry electrodes | 12, 19, 19 | accuracy, FOR (false omssion rate), sensitivity, SNR | Matlab, EEGlab | - | Findings: (1) New fingerprint method based on temporal, spectral and spatial methods, that are used together as an input for SVM learning strategy; (2) Comparison of dry and gel based systems; (3) Combination of model with components originating from multiple artifact sources at the same time (cardiac, myogenic, eyes, ...) |  | Very informative paper with a very good introduction.
Lin18 | 281 | Real-Time EEG Signal Enhancement Using Canonical Correlation Analysis and Gaussian Mixture Clustering | Lin CT, Huang CS, Yang WY, Singh AK, Chuang CH, Wang YK. | J Healthc Eng | cs | attention strategy | The subjects were stimulated with VEP and SSVEP while no motion, blinking, chewing and head rotation. The evoked state had to be classified | EEG cap with 62 Ag/AgCl electrodes, with two references (opposite lateral mastoids). Consistent with the modified international 10-20 system. Impedance below 5kOhm, 1kHz, Scan SynAmps2 Express System | 11 right-handed (7 males) between 18 and 27 years. No known psychological disorders | 11 | visual (fit of corrected EEG-signals) | - | no | Combination of CCA and Gaussian model for artifact removal |  | Weak result and no real evaluation metric, but very informative
Delisle-Rodriguez17 | 293 | Adaptive Spatial Filter Based on Similarity Indices to Preserve the Neural Information on EEG Signals during On-Line Processing | Delisle-Rodriguez D, Villa-Parra AC, Bastos-Filho T, López-Delis A, Frizera-Neto A, Krishnan S, Rocon E. | Sensors (Basel) | cs | SSVEP | The subjects (1) made a SSVEP task with 40 characters of frequencies 8-15.8Hz that are classified and (2) a BCI experiment for gait planning recognition | (1) Synamps2 (Neuroscan) with 64-channels in the 10-20 system, sampling rate of 1000Hz and downsampled to 250Hz and (2) BrainNet BNT 36 (EMSA, Brazil) with 20 selected channels from 64 gel Ag/AgCl electrodes for the BCI experiment | (1) 35 subjects (17 females, aged 17-34 years), (2) 6 healthy male subjects (mean age 31 years) | 35, 6 |  | - | - | Adaptive Filtering for LAR and WAR improve the results. The features fractal dimension, reference-free EEG, and power density spectra from 0.1-30Hz are important (but others say >30 is important for walking) |  | Complicated to read and not very clear
Wang17 | 300 | Stretchable Dry Electrodes with Concentric Ring Geometry for Enhancing Spatial Resolution in Electrophysiology | Wang K, Parekh U, Pailla T, Garudadri H, Gilja V, Ng TN. | Adv Healthc Mater | cs | - | The subject was asked to alternatively open and close his eyes to see the alpha-band activities | Proposed strechable and printed electrodes with was sticked onto his forehead based at the F3 position of the 10-20 system. OpenBCI board set to amplify, 250 Hz, Reference at the A1 and A2 (ear lobes) | One male | 1 | visual | Matlab | no | Proposed dry concentric electrodes printed on a strechable layer, that are able to perform a Laplacian filter. This removes the noise effectively, but can only measure near sources. |  | many dry electrodes could be good for us.
Jas17 | 301 | Autoreject: Automated artifact rejection for MEG and EEG data | Jas M, Engemann DA, Bekhti Y, Raimondo F, Gramfort A. | Neuroimage | mst | AEP, VEP, motor imagery, attention strategy | Proposes fully automated method "Autoreject". Method is compared to SNS, FASTER and RANSAC in 4 open data sets with over 200 participants used to evaluate the methods | Paper worked with 4 different databases, used 60-64 channel data subjects and in one case 248 magnetometers MNE sample data - Neuromag VectorView system + MEG-compatible cap with 60 electrodes was used for the EEG recordings (sampled at 150 Hz) EEGBCI dataset - 64-channel BCI2000 EEG cap EEG faces data (OpenfMRI ds000117) - 70-channel Easycap EEG Human Connectome Project (HCP) MEG data - multiple different devices used | validated methods on 4 open datasets with data from over 200 subjects MNE sample data - multimodal open dataset consisting of MEG and EEG data. Number of subjects total unclear. EEGBCI dataset - motor-Imagery dataset with 109 subjects EEG faces data (OpenfMRI ds000117) - multimodal task-related neuroimaging data for EEG, MEG and fMRI. Number of subjects total unclear. Human Connectome Project (HCP) MEG data - multimodal dataset by multiple international laboratories around the world. It provides task-free and task-related data from over 900 subjects | 200 | infinity norm of evoked responses between method estimate and ground truth | Matlab, Python | yes | Presented a novel algorithm for automatic data-driven detection and repair of bad segments in single trial M/EEG data. Validation suggests that autoreject performs at least as good as diverse alternatives and commonly used procedures. This is the consequence of the combination of a data-driven outlier-detection step combined with physics-driven channel repair where all parameters are calibrated using a cross-validation strategy robust to outliers. An underlying assumption however is that the artifacts are deviations of the signals far from zero. | Proposes automatic and data-driven algorithmic solution to a task normally done manually. Autoreject reduces the cost of data inspection by experts. A comparison to visual inspection by experts therefore would be intersting to see | Autoreject Sources are available here: https://autoreject.github.io
Oliveira17 | 328 | A Channel Rejection Method for Attenuating Motion-Related Artifacts in EEG Recordings during Walking | Oliveira AS, Schlink BR, Hairston WD, König P, Ferris DP. | Front Neurosci | cs | attention strategy | The subjects walked on a treadmill alternated with 1) open and 2) closed eyes for 5 times 5 minutes each. The proposed method should remove motion artifacts such that the signals are independent of the gait cycle. | The subjects walked on a treadmill alternated with 1) open and 2) closed eyes for 5 times 5 minutes each. The proposed method should remove motion artifacts such that the signals are independent of the gait cycle. | 10 subjects (6 males) between 21 and 36 years. No known locomotor deficits | 10 | PDS (spectral power) | Matlab | no | They found that there are ~5-30 less independent components if the subjects walk blindfolded. Higher number of EEG channels improve the results (may due to ICA) | Experiments with varying walking and running speeds | References quite often Gwin10, the proposed method is not good
DelPozo-Banos17 | 333 | Localized component filtering for electroencephalogram artifact rejection | DelPozo-Baños M, Weidemann CT. | Psychophysiology | mst | attention strategy | proposes new post-step methodology to enhance the performance of existing BSS systems: Localized component filtering (LCF) | 133 channels (128 scalp channels, 2 mastoid channels, and 3 EOG channels) using BioSemi Active Two system and a sampling rate of 500Hz | real data set recorded from 18 participants while performing a recognition memory task | 18 | true acceptance rate (TAR), false acceptance rate (FAR) | Python | yes | method takes the original components from BSS as inputs, and mixes them so that the processed (cleaned) components replace the original ones only when an artifact is detected. | the proposed LCF method should be validated against multiple pipelines of this list | LCF method can be included into any other existing pipeline that has been added here, not sure if it actually works
Weiss17 | 345 | AR2, a novel automatic muscle artifact reduction software method for ictal EEG interpretation: Validation and comparison of performance with commercially available software | Weiss SA, Asadi-Pooya AA, Vangala S, Moy S, Wyeth DH, Orosz I, Gibbs M, Schrader L, Lerner J, Cheng CK, Chang E, Rajaraman R, Keselman I, Churchman P, Bower-Baca C, Numis AL, Ho MG, Rao L, Bhat A, Suski J, Asadollahi M, Ambrose T, Fernandez A, Nei M, Skidmore C, Mintzer S, Eliashiv DS, Mathern GW, Nuwer MR, Sperling M, Engel J Jr, Stern JM. | F1000Res | mst | - | Evaluates newly proposed method for EMG contamination reduction (AR2) against existing "Persyst v12 artifact reduction software" (AR1) in a blinded expert review. Goal is to detect seizures. New method is based on CUDAICA. | EEGs were acquired using "EEG-1200" amplifier (Nihon-Kohden, Irvine, CA) at sampling rate 200 Hz and a low frequency cut-off of 0.08 Hz | 8 adult patients with suspected focal-onset seizures were selected based on a review of consecutive clinical neurophysiology case conference presentations (between January 1, 2014 and December 1, 2015). The conference consensus was that these ictal EEG records were uninterpretable due to muscle artifact contamination when reviewed with conventional digital filtering | 8 | visual inspection | Matlab, Persyst(v12) | yes | A new open-code (github) implementation for CUDAICA-based artifact removal is introduced and compared against existing methods from Persystv12 implementation in seizure detection use-cases | A comparison of results here was not possible due to Persystv12 methods not being described in detail. Further work suggested: evaluating proposed CUDAICA against other ICA implementations from the list here. CUDAICA might be especially interesting for its expecteded speed performance | CUDAICA implementation for AR2 can be found here: https://github.com/shennanw/AR2
Roy17 | 347 | Gaussian Elimination-Based Novel Canonical Correlation Analysis Method for EEG Motion Artifact Removal | Roy V, Shukla S, Shukla PK, Rawat P. | J Healthc Eng | cs | - | Proposed GECCA with improved | 24-channel EEG with 256 Hz sample rate for 10 seconds, the number of channels was reduced to 16 | A dataset from PhysioNet was used |  | RMSE, correlation, lambda, SNR | - | no | Improved the performance of CCA using Gaussian Elimination (GECCA) by applying it on IMFs from an EEMD and subsequently a SWT. GECCA is 18% faster, the accuracy should also be better, but the ROC shows that this depends on the sensitivity. SWT is more stable than DWT. |  | The superiority of GECCA is not that clear, as it depends on the used sensitivity in the ROC analysis. They cite that CCA is more accurate and faster than ICA, especially for muscle artifacts [Soomro14].
Melman16 | 370 | Robust power spectral estimation for EEG data | Melman T, Victor JD. | J Neurosci Methods | cs | - | Altered a tampered based power density estimator to work with quantile-based aggregations like the median | 128-channel EEG on a 10-20 system, sampled at 250Hz | One male subject with 23 years | 23 | PDS (power density spectrum) | Matlab, EEGlab | yes | The estimation of quantiles (median) in the power density estimation is superior for any non-gaussian distribution |  | Very interesting method of calculating a robust PDS estimation for non-gaussian distributions
Oliveira16 | 377 | Proposing Metrics for Benchmarking Novel EEG Technologies Towards Real-World Measurements | Oliveira AS, Schlink BR, Hairston WD, König P, Ferris DP. | Front Hum Neurosci | cs | attention strategy | For each of the three setups, four blocks of 160 trials of an auditory oddball experiment. Different evaluation metrics are applied for the different condition (seated, walking) and EEG-system | 256-channel EEG (BioSemi ActiveTwo), 72-channel Cognitiontics Mobile (gel) and Cognitiotics Dry (dry), each sampled at 500Hz | 7 healthy participants (4 males) | 7 | pre-stimulus noise (PSN), SNR, variance accross P300 event window (CV), scalar products | Matlab, EEGlab | - | They found that dry electrodes were not as successfull as gel electrodes and the noise was higher during walking | - | Performing a manual channel rejection in advance, Useless metrics for evaluating if methods retain brain signal that can be used for a discrimination or estimation task, Only ICA is used and we see multiple characteristics from it
Cheveigne16 | 389 | Sparse time artifact removal | de Cheveigné A. | J Neurosci Methods | cs | - | The subjects were measured without any specific objective | simulated data (1) and a 128-channel EEGsystem (Biosemi). The data sampled at 512 Hz | 8 subjects measured about 1.5–2.5 h each | 8 | visual fit (of corrected EEG-signals) | Matlab | yes | novel STAR-algorithm that removes channel-specific noise from electrodes, sensors and certain forms of myogenic artifacts. The time axis is partitioned into an artifact-free and an artifact-contaminated part, and the correlation structure of the data is estimated from the covariance matrix of the Artifact-free part. Artifacts are then corrected by projection of each channel onto the subspace spanned by the other channels. Kurtosis and channel-independence is an indicator for artifacts. Component Analysis can't handle more dimensions than channels |  | Does not work if an artifact affects two or multiple sensors. STAR can be used complementary
Costa16 | 392 | Characterization of Artifacts Produced by Gel Displacement on Non-invasive Brain-Machine Interfaces during Ambulation | Costa Á, Salazar-Varas R, Úbeda A, Azorín JM. | Front Neurosci | cs | -, motor imagery tasks | Characterization of gel displacement on four datasets | gel electrodes according to the international 10/10 system with 1200Hz sample rate. 32 electrodes fro the first 3 datasets and 16 for the last one. | Four datasets, three without mental strategy and one motor imagery task. 3 to 12 subjects of mean age between 25 and 30. | 3, 12 | PDS, SMA (Segments Maximum Average) | - | no | Sudden amplitude changes happen often in moving setups, but hardly in movement-free. High amplitude changes are independent of the movements
Nathan16 | 400 | Negligible Motion Artifacts in Scalp Electroencephalography (EEG) During Treadmill Walking | Nathan K, Contreras-Vidal JL. | Front Hum Neurosci | cs | motor execution | Characterization of EEG signal for different walking speeds (1.5, 3.0, 4.5km/h) and filtering with Fourier, wavelet and ASR | 64-channel active electrodes (BrainAmpDC, actiCap) sampled at 1000Hz, additional IMU with 128Hz | 4 able-bodies (3 males) aged 26-33 years | 4 | visual, phase lag between EEG and accelerometer | Matlab | - | Group-averaged data showed that the acceleration is leading the EEG signal in the phase. This is also clear, as the muscles as recognised in the EEG counterbalances the head. |  | Averaging over multiple gaits, trials and even subjects
Zou16 | 405 | Automatic Identification of Artifact-Related Independent Components for Artifact Removal in EEG Recordings | Zou Y, Nathan V, Jafari R. | IEEE J Biomed Health Inform | cs | attention strategy | The subjects did a P300 speller task in which they focused visually on letters within an alphabet. 2-5 sessions of 20-30 letters. | A custom platform based on a TI ADS1299 with 8 wet electrodes, aditionally a TI ADS1299 that produces currents of 24nA to measure the electrodes impedances | 10 healthy subjects with no previous experience with the P300 speller taks | 10 | accuracy (of artifactual epochs compared to experts) | Matlab | - | They use a current generator with 24nA and 30.5Hz to measure the impedance of each electrode-scalp contact. Then they used the unmixing matrix from the previous ICA and the impedance to weight the electrode for the back-projection. | - | The result should be weighted weak, as they perform an ICA on only 8 channels and compare their method with other approaches that are partly not built for non-biological artifacts. Applies harsh averarging over multiple subjects, trials, gaits, etc.
Gordon15 | 406 | Informed decomposition of electroencephalographic data | Gordon SM, Lawhern V, Passaro AD, McDowell K. | J Neurosci Methods | mst | attention strategy, motor execution, motor tapping study | New approach "Informed Multidimensional ICA (IMICA)" that builds upon the constrained optimization approaches for ICA for better model and subspaces construction. Uses likelihood function to find the optimal model size for each target subspace. | First data set: 64 channel BioSemi Active II system + eye tracking with FaceLAB eye tracking system Second data set: 256 channel BioSemi Active II system | First data set: 10 right-handed, male subjects (age 20–40) Second data set: 8 right-handed males and 2right-handed females | 10, 8 | Standard Deviation, Correlation, Components removed, visual | - | - | Previously existing decomposition methods tend to suffer from the limitation of being inherently blind and therefore cannot leverage domain-specific knowledge to improve the overall quality of results. IMICA attempts to fill this gap | - | -
Snyder15 | 408 | Independent Component Analysis of Gait-Related Movement Artifact Recorded using EEG Electrodes during Treadmill Walking | It is problematic that the impedance of the electrodes is very high and therefore only simulated artifacts from loose electrodes. The impedance of their setup is a factor 10 to 1000 higher as from a real setup (900kOhm). Nonetheless, this is neccessary. | Front Hum Neurosci | mst | motor execution | Separating true brain signals from motion artifacts by blocking neuronal activity with a novel setup using a silicon cap and artificial scalp while different walking speeds and EEG data collected | 256-channel active electrode array system (BioSemi Active 2); recorded movement artifact at 512 Hz | 10 healthy young participants | 10 | mutual information, correlation coefficient | Matlab | no | ICA and dipole fitting analysis on pure gait-related movement artifacts can offer some insights into how to interpret ICAs performed on walking data. The results indicate that components can be found in cortical areas where activity is expected during walking, even with pure artifact data as an input, and that for about half the subjects, these components can be reliably found across different data sets. | Alternative methods are needed for the identification and separation of movement artifact in mobile EEG signals, especially methods that can be performed in real time. | Offers some insights in researching movement artifacts separately from neuronal activity; performance metrics focuses on the channel variance of EEG channels and the correlation to neighboring chnnels; ground truth has been established with ground reaction force treadmill
Kline15 | 420 | Isolating gait-related movement artifacts in electroencephalography during human walking | Kline JE, Huang HJ, Snyder KL, Ferris DP. | J Neural Eng | cs | motor execution | Subjects wore a swimming cap that isolates the electrodes and they evaluated which methods could reduce the artifacts from the cable swings | 256-channel active electrode array system (BioSemi Active 2); recorded movement artifact at 512 Hz | 9 healthy young participants (4 males), mean age 27 | 9 | visual | Matlab | - |  |  | It is problematic that the impedance of the electrodes is very high and therefore only simulated artifacts from loose electrodes. The impedance of their setup is a factor 10 to 1000 higher as from a real setup (900kOhm). Nonetheless, this is neccessary
Raduntz15 | 431 | EEG artifact elimination by extraction of ICA-component features using image processing algorithms | Radüntz T, Scouten J, Hochmuth O, Meffert B. | J Neurosci Methods | mst | nine different cognitive tasks (e.g. attention strategy, motor execution, ...) | Proposes automated artifact removal using linear discriminant analysis (LDA) for classification of feature vectors extracted from ICA components via image processing algorithms. Results compared to expert visual classification | 25 electrodes arranged according to the 10–20 system (amplifier from Brain-Products GmbH and corresponding BrainRecorder software.) | 57 participants (31 female, 26 male, age: 30-62). Sample shows high variability with respect to cognitive capacity and age. | 57 | accuracy, SNR, performance time | Matlab, BrainRecorder software, E-Prime application suite | - | Describes new automated approach (based on machine classification of ICs) for EEG artifact rejection inspired by computer vision. Features obtained with image processing algorithms. The new part really seems to be the feature extraction through LBP + gradient images + range filter + laplacian. | Feature extraction through image processing could be compared to traditional Feature extraction in EEG | -
Mariani15 | 444 | Use of multiscale entropy to facilitate artifact detection in electroencephalographic signals | Mariani S, Borges AF, Henriques T, Goldberger AL, Costa MD. | Annu Int Conf IEEE Eng Med Biol Soc | cs | - | Using a method based on multiscale Entropy to detect artifactual segments in EEG windows | 2 EEG signals from the prefrontal cortex at 2048Hz sample rate, additional 3-axis accelerometer with 200Hz | 23 recordings with 8-9 minutes each | 23 | AUC | - | - | The CI-metric that is based on multiscale entropy is well suited to detect artifactual components in comparision with standard deviation. | Explore more than one type of artifact, apply an ensemble of detectors to different cases. Further studies to assess the utility and limitations in comparison with currently used techniques. | The AUC measure of the ROC curve is very well for classificators that could be applied in multiple use cases.
Zhang15 | 447 | Automatic Artifact Removal from Electroencephalogram Data Based on A Priori Artifact Information | Zhang C, Tong L, Zeng Y, Jiang J, Bu H, Yan B, Li J. | Biomed Res Int | cs | motor imagery, attention strategy | Subjects were given motor imagery and visual stimuli. The pipeline was used to classify the type of the input. A session to measure the individual artifacts allone was performed in multiple trials and for multiple movements. | (1) EEG 16-channel EEG (gtec), subset of the 10-20 system with Ag/AgCl electrodes with a sample rate of 512Hz. (2) EEG same setup | (1)14 healthy participants. (2) 13 healthy participants aged 20 to 24 years. | 14, 13 | accuracy | Matlab | - | They found that a wICA is better than doing nothing. This works much better for EOG than for EMG artifacts. |  | Not that relevant for us.
Bulea14 | 451 | Sitting and standing intention can be decoded from scalp EEG recorded prior to movement execution | Bulea TC, Prasad S, Kilicarslan A, Contreras-Vidal JL. | Front Neurosci | cs | motor execution | Subjects were told to (1) self-intented or (2) auditory triggered to sit down and stand up. Pre-Movement EEG was used to classify the stages (1) quite, (2) sit, or (3) stand in each 10 alternating sit-down and stand-up transitions with a rest between 3 and 10 seconds. | 64-channel active electrode (Brain Products) at 1000Hz with the 10-20 system and impedance below 25kOhm | 10 healthy adults (6 males) | 10 | accuracy | Matlab | no | Features from the theta band are well suited to classify movements from pre-motion EEG both in the self-paced and triggered case. GMM needs only one component per class for most subjects. | Reduce the number of freedom in the parameter optimization. Include a learning phase of the hardware to increase the accuracy. Include EMG data to increase the accuracy. | Very strong focus on little delay of the prediction, good references of how to use pre-motion signals in BMI, indicate that EEG+IMU might be good. No real artifact correction or removal
Chen14 | 455 | A preliminary study of muscular artifact cancellation in single-channel EEG | Chen X, Liu A, Peng H, Ward RK. | Sensors (Basel) | mst | motor execution | Proposes to combine ensemble empirical mode decomposition (EEMD) and joint blind source separation (JBSS) techniques for single channel artifact removal. Compares the method to other existing single channel artifact removal methods | SynAmps2 amplifiers (NeuroScan, Compumedics, El Paso, TX, USA) at 1000Hz sample rate | Dataset_1: ictal activity from a patient with mesial temporal lobe epilepsy Dataset_2: eight healthy subjects while stably cycling on an exercise bicycle | 1, 8 | SNR | Matlab | - | Claims that SCICA and EEMD-ICA are unsuitable for muscular artifact cancellation in the single-channel EEG case, because of ICA permutation problem. In there dataset there was remaining brain activity AND muscle activity in almost all components when the two methods above were used. In contrast, EEMD-CCA and EEMD-IVA were able to effectively isolate muscle activity into the bottom components due to their low autocorrelations. | This single channel comparison study is now many years old and could benefit from another comparison study including newer single channel methods | Is a good first read as an entry point into single channel artifact removal as it explains many of the basics
Yilmaz14 | 465 | Interference of tonic muscle activity on the EEG: a single motor unit study | Yilmaz G, Ungan P, Sebik O, Uginčius P, Türker KS. | Front Hum Neurosci | mst | motor execution | Evaluated the impact of the temporalis muscle (single motor unit - one individual muscle contribution) on the EEG. The muscle also impacts the EEG in resting state (!!!) which makes it especialy interesting for all EEG artifact removal pipelines. Evaluated with three resting states: eyes open, eyes closed, jaw dropped. | EEG activity recorded with 10–20 system (MEDCAP , SpesMedica S.r.l, Italy) with 21 Ag/AgCl electrodes (Fp1, Fp2, Fz, F3, F4, F7, F8, C3, C4, T3, T4, T5, T6, Pz, P3, P4, O1, O2, all referenced to Cz). Used conductive gel (Electro-Gel, ECI, USA). SMU activity was recorded using 2 intramuscular silver fine-wire electrodes coated with Teflon (75μm in core diameter; Medwire, USA). | Data collected from 9 healthy male volunteers (aged 20 – 42 years) | 9 | visual | -> SystemPlus(Micromed) for recording -> Spike2 (Cambridge Electronic Design) for analysis | yes | Did not find significant difference between the eyes open and closed states regarding the interference from the SMUs, but it wasevident that some SMUs were silenced in the eyes closed condition. Therefore recommends that the EEG should be recorded in eyes closed and jaw dropped conditions to decrease the amount of contamination from the resting activity of the jaw muscles and to generate a genuine relaxation of the subject. | Single motor unit studies with an attempted ground truth measurment have a very high value as they add knowledge of individual muscle contribution to the EEG system. There should be more studies targeting different muscles or muslce groups, but they seem to be rare. | Very interesting paper to understand the individual EEG contamination contribution from the temporalis muscle. We need more studies like this for different muscles or muscle groups in the future!!!
Reis14 | 482 | Methodological aspects of EEG and body dynamics measurements during motion | Reis PM, Hebenstreit F, Gabsteiger F, von Tscharner V, Lochmann M. | Front Hum Neurosci | cs | - | Comparision of hardware, software and techniques for EEG with high motion aspect in form of secondary literature. | different specialized systems with the goal to work mobile | - |  | - |  | no | AMICA is better than ICA. ACEMIC and MARA find IC with motion artifacts automatically. Iterative IC rejection is sucessful |  | Handles electrodes, caps, algorithms, frameworks and methods to get a clean EEG while moving. Can be regarded as survey.
Boudet14 | 490 | Dual adaptive filtering by optimal projection applied to filter muscle artifacts on EEG and comparative study | Boudet S, Peyrodie L, Szurhaj W, Bolo N, Pinti A, Gallois P. | ScientificWorldJournal | mst | - | Proposes new two-step method "dual adaptive filtering by optimal projection (DAFOP)" to automatically remove artifacts while preserving true cerebral signals. The method uses common spatial patterns (CSP) and a regression method for signal reconstruction. Results are compared to existing BSS-CCA method and standard filtering methods. | 19 channels according to the 10/20 system. Clinical EEG recordings evaluated by two neurologists | Clinical recordings of 20 epileptic patients with pertinent cerebral rhythms were selected. Among these 20 recordings, 114 relevant pages (1 page = 20 s epoch of EEG) with different artifact contamination were selected. Previous training was done with a different dataset. | 20 | Percentage of estimated removed artifacts, accuracy, visual | http://www.neurology-kuleuven.be/ | - | DAFOP method is being introduced | Method should be evaluated against stronger CSP alternatives, rather than just BSS-CCA and Filtering | Weak comparison. Why CCA and Filtering?
Li13 | 509 | L1 norm based common spatial patterns decomposition for scalp EEG BCI | Li P, Xu P, Zhang R, Guo L, Yao D. | Biomed Eng Online | mst | motor imagery | Paper presents L1 norm based CSP (common spatial pattern) implementation and compares it to existing conventional L2 norm CPS and regularized CSP with Tikhonov regularization (TR-CSP) | 15 Ag/AgCl electrodes covers sensorimotor area were used to record the EEG, and the signals were sampled with 1000 Hz | EEG data from 13 subjects. During online experiment, subjects were required to sit in a comfortable armchair in front of a computer screen, and they were asked to perform motor imagery with left hand or right hand according to the instructions appeared on the screen. Motor imagery lasts for 5 seconds, and follows a 5 seconds rest. | 13 | accuracy (for classification) | - | - | shows that L1-CSP is feasible | probably benefits from neutral comparison study by other authors comparing multiple CSP versions | Seems to be the only paper using common spatial patterns we found in our search, which is strange. Also one of the rare occasions we have a paradigm with motor imagery tasks. Paper is only weakly tied to motion artifacts through head movement
Lawhern13 | 524 | DETECT: a MATLAB toolbox for event detection and identification in time series, with applications to artifact detection in EEG signals | Lawhern V, Hairston WD, Robbins K. | PLoS One | mst | attention strategy | First introduction of the MATLAB toolbox DETECT, which can be used for automatic training and identification of events | EEG recorded at 512 Hz using a 64-channel Biosemi ActiveTwo system (Biosemi, Amsterdam, NL) and referenced to the average of the 2 mastoids. 4 external channels used to record eye movements by electrooculography (EOG). | Data was acquired from three participants who performed a standard visual-evoked-potential (VEP) task, in which they reported whether each presentation of an image was a depiction of a U.S. Soldier (common) or an enemy insurgent (infrequent | 3 | visual | Matlab, DETECT | yes | Introduces DETECT toolbox and shows its application for studies with multiple different sources of events with artifacts (e.g. muscle, eye activity and heart rate). NO comparison is done. | - | -
Nottage13 | 538 | A novel method for reducing the effect of tonic muscle activity on the gamma band of the scalp EEG | Nottage JF, Morrison PD, Williams SC, Ffytche DH. | Brain Topogr | mst | motor execution | Tonic Muscle Artefact Reduction in Gamma Band | Recorded with Neuroscan software and a 64 channel cap using the 10–20 system. Additionally 2 electrodes to record left and right eye saccade (Saccade Muscle Potential = SMP). Recorded Group 1 with 2000Hz and Group 2 with 5000 Hz | 2 independent groups, of 8 subjects each (16 total); using two different EEG recording protocols; both groups performed self-paced motor task (button pressing 70 times); 5 females and 3 males in group 1 and also in group 2 | 16 | SNR | Matlab | - | Reduction of EMG contamination achieved by fitting and subtraction of individual scalp muscle spikes using Gabor functions leads to improvement of the signal to noise ratio. The paper also reports multiple observed effects of muscle contamination on the gamma frequency band | Paper claims that fitting of individual muscle spikes is promising, but this work was the first to use it. It should be further investigated if this method works in different pipelines and/or on other frequency bands | Might be interesting to look at gamma band for artifact detection but correction still needs to be done in alpha/beta band for our work
Yong12 | 549 | Automatic artefact removal in a self-paced hybrid brain- computer interface system | Yong X, Fatourechi M, Ward RK, Birch GE. | J Neuroeng Rehabil | cs | motor execution | Participants wrote sentences by focusing on visual alphabet and attempted hand extension. They investigated different wavelet and BSS strategies to remove artificial noise from the semi-simulated EEG | Hybrid system of eye-tracker (Mirametrix S1) and 64-channel electrode cap. They used 15 channels on the motor cortex area, 4 for EOG and 8 for EMG for the facial muscles. Sampled at 128Hz with L64 Sagura. | 7 able-bodied subjects of age 26 to 31 | 7 | MSE, Spectral distortion | Matlab |  | Wavelet filters with adaptive thresholds are better suited than hard thresholds and also SOBI, ERICA and AMUSE. Especially ASWTD with Garrote thresholding function exceeds. Hard thresholding doesn't work well because of the incontinuous signal | Implement in C++ to increase the speed even more (now 200ms latency). Automatically select the optimal wavelet function. Extend from the univariate version to a multivariate version. | Very good introduction and sources for the challenges and artifact strategies for BCI systems. Emphasize the importance of idle states to reduce False Positives in the most times. Unfortunately only tested for simulated artifacts
Lau12 | 550 | Weighted phase lag index stability as an artifact resistant measure to detect cognitive EEG activity during locomotion | Lau TM, Gwin JT, McDowell KG, Ferris DP. | J Neuroeng Rehabil | cs | attention strategy | Participants were standing or walking (0.8m/s) while attending a visual oddball experiment. A newly proposed metric weighted phase lag index is evaluated for a measure of artifact restistance | 248-channel EEG ActiveTwo (BioSemi) sampled at 512Hz | 8 helthy participants of age 20-31 years | 8 | visual | Matlab | - | The results indicate that the proposed WPLI metric is can reduce non-brain signals efficiently. The results show that the P300 peak has a high inter- and intra-subject variance, indicating that this is not a good measure for the oddball experiment |  | Very interesting if it is further worked on. Subresults of this method could be used as additional input for a larger network
Lawhern12 | 551 | Detection and classification of subject-generated artifacts in EEG signals using autoregressive models | Lawhern V, Hairston WD, McDowell K, Westerfield M, Robbins K. | J Neurosci Methods | cs | motor execution | Subjects were told to make one of different movements. They used autoregression methods in order to classify the type of movement artifact in the EEG. | 64-channel EEG (Biosemi ActiveTwo) at a sample rate of 512Hz, downsampled to 256Hz. Additional 4 EOG electrodes to verify the blinks, not for the processing. | 7 participants. | 7 | Accuracy | Matlab, EEGLAB, ERPLAB | - | The combination of auto-regressive models and SVM that use the parameters of the first as input is successful to discriminate the type of artifact. | Build an online tool that selects artifact-free epochs. Use multi-variate auto-regressive (MVAR) methods. Test the model for stability.
Winkler11 | 583 | Automatic classification of artifactual ICA-components for artifact removal in EEG signals | Winkler I, Haufe S, Tangermann M. | Behav Brain Funct | mst | motor imagery, attention strategy | Proposes an automatic method to classify general artifact ICs based on TDSEP (Temporal Decorrelation source SEParation) and LPM (Linear Programming Machine), LDA (Linear Disc. Analysis) and CSP (Common Spatial Patterns) | Setup 1: 121 electrodes used, no further information provided Setup 2: 64 Ag/AgCl electrodes, no further information provided Setup 3: 90 EEG channels, no further information provided | Setup 1: Data from 12 healthy right-handed male subjects Setup 2: A group of 18 subjects of 20 to 57 years of age (mean =34.1, SD = 11.4) Setup 3: 80 healthy BCI-novices performed first motor imagery with the left hand, right hand and both feet in a calibration (i.e. without feedback) measurement. | 12,18,80 | accuracy (for classification) | Matlab | yes | Optimized linear classifier only needed 6 features to perform on level with the inter-expert disagreement (<10% Mean Squared Error (MSE)) | [1] suggests to test the used classifiers performance with reduced electrode sets [2] suggests to investigate wether correctly detected artifacts can even always lead to better motor imagery classification accuracy | Has a very high amount of participants, how does that even work out?
Gwin10 | 615 | Removal of movement artifact from high-density EEG recorded during walking and running | Gwin JT, Gramann K, Makeig S, Ferris DP. | J Neurophysiol | cs | attention strategy | visual oddball discrimination task performed while standing, walking (0.8 and 1.25m/s) and running (1.9 m/s). Two pipelines were evaluated how good they removed motion artifacts | ActiveTwo amplifier with a 248-channel EEG (BioSemi) and gel electrodes, 512Hz | 8 healthy persons (7 males, 21-31 years old) | 8 | mean spectral power (within 1.5-8.5Hz) | Matlab, DataSuite | no | First paper on how to reduce motion artifacts in walking and running. The proposed two algorithms that use linear fits and AMICA. | Improve pipeline, record head kinematics to make usable in real-time, use mobile EEG systems that are in development to reduce cable swing artifacts | Goal is very close to ours, multiple good references included
McMenamin10 | 624 | Validation of ICA-based myogenic artifact correction for scalp and source-localized EEG | McMenamin BW, Shackman AJ, Maxwell JS, Bachhuber DR, Koppenhaver AM, Greischar LL, Davidson RJ. | Neuroimage | cs | motor execution | Validation of ICA for myogenic artifact correction and source localization on artificially contaminated EEG data in a setting of 4 conditions: eyes open/closed, muscles tense/relaxed. | Geodesic Sensor Net with 128 electrodes, sampled at 500Hz, bad channels and gross artifacts were rejected manually. 16 blocks of 32 seconds of the same conditions. | 17 persons (16 female, mean age 24.1 years) | 17 | sensitivity, specificity, | Matlab | no |  | Validate during a neuromuscular blockade to remove EMG totally. Comparision of more methods. | Pretty poor specificity, maybe because they used a PCA prior to ICA
Mosher09 | 640 | GENERALIZED SIDELOBE CANCELLER FOR MAGNETOENCEPHALOGRAPHY ARRAYS | Mosher JC, Hämäläinen MS, Pantazis D, Hui HB, Burgess RC, Leahy RM. | Proc IEEE Int Symp Biomed Imaging | mst | attention strategy | Introduction of the the "generalized sidelobe canceller (GSC)" method as an alternative to the already existing "linearly constrained minimum variance (LCMV)" algorithm. Method was applied with MEG device, but works for EEG as well | Data acquired with MEG instrument with 102 magnetometers and 204 planar gradiometers (recorded at 1000 samples per second). No testing of EEG instrument was done in this paper!! | One human subject which had lightly magnetized implants and dental work that created additional artifacts inside the shielded room. The subject’s left and right median nerves were stimulated in a standard somatosensory evoked field protocol, for 266 repetitions for each hand over a total of about 200 seconds | 1 | visual | - | - | Introduction of beamformer method "generalized sidelobe canceller (GSC)" | [1] Needs study to investigate wether the proposed GSC method can be used for muscle artifacts in EEG recordings [2] Needs comparison study to evaluate wether the proposed GSC method performs better than the alternative LCMV approach | Note: No testing of EEG instrument was done in this paper!! Needs evaluation!
McMenamin09 | 648 | Validation of regression-based myogenic correction techniques for scalp and source-localized EEG | McMenamin BW, Shackman AJ, Maxwell JS, Greischar LL, Davidson RJ. | Psychophysiology | cs | motor execution | Participants were asked to open/close their eyes and relax/tension their facial muscles. They investigated designs to localize neural activity | Sixteen blocks of 32-second trials with different cases. 128-channel Geodesic Sensor net, sampled at 500Hz. | 17 subjects (16 females, mean age 24.1 years) | 17 | sensitivity, specificity | Matlab | no | LORETA is not well suited to correct EMG artifacts | Usage of ICA
Delorme04 | 707 | EEGLAB: an open source toolbox for analysis of single-trial EEG dynamics including independent component analysis | Delorme A, Makeig S. | J Neurosci Methods | mst | attention strategy, motor execution | First introduction of the EEGLAB as a toolbox and graphic user interface running under MATLAB. EEGLAB functions include EEG data import and visualization, preprocessing (including artifact rejection, filtering, epoch selection, and averaging), Independent Component Analysis and time/frequency decompositions. | The test sample dataset consists of 80 3-second EEG epochs time-locked to targets presented in the left visual field between 3 and 1.5 degrees of visual angle. Data from thirty-one scalp electrodes (referred to the right mastoid) were sampled at 500 Hz (later reduced for compactness to 125 Hz). Dataset also available from http://www.sccn.ucsd.edu/eeglab/ | Paper references an experiment in which one subject covertly attended a cued location on the computer screen, responding quickly with a thumb button press each time a target (filled square) was briefly presented at this location (Makeig et al., 1999). In different trial blocks, the attended location was any one of five positions arranged horizontally on the computer screen above a fixation cross. Unfortunately no additional information is revealed about the subject. | 1 | visual | Matlab, EEGLAB | yes | The paper primarily showcases the functions of EEGLAB. A small dataset was provided to test the functions. | Authors discuss implementing Hilbert transform in addition to existing linear filters and that there still is a lack in assessing bi-coherence between frequencies in EEGLAB | Has been used a lot by many researchers eversince, so there already is a lot of follow-up research present
Iriarte03 | 713 | Independent component analysis as a tool to eliminate artifacts in EEG: a quantitative study | Iriarte J, Urrestarazu E, Valencia M, Alegre M, Malanda A, Viteri C, Artieda J. | J Clin Neurophysiol | mst | - | Objectively test the capability of ICA (JADE algorithm) to eliminate different artifacts in recordings without altering the rest of the signal by comparing artifact-free and cleaned fragments of EEG. Goal was to prove that ICA removes some components, but does not modify the rest | 64-channel digital EEG using LaMont amplifiers (LaMont Medical, Madison, WI, U.S.A) and Harmonie software (Stellate, Montreal, Quebec, Canada) | Two series of samples of interictal EEG (between 4 and 10 seconds each) were selected from two different patients. 40 EEG artifact fragments (with at least one spike, eye movement, EKG, line noise, muscle activity and electrode artifact) selected per subject. 10 artifact-free fragments selected per subject. Patient 1 had symptomatic frontal epilepsy. Patient 2 had focal epilepsy caused by unilateral mesiotemporal sclerosis | 2 | Visual | Matlab | - | Demonstrated/Validated utility of ICA (JADE algorithm) for visual and quantitative usage. Also mentions some problems and limitations when working with ICA: (1) lengths of the sample (2) exact number of components (3) selection of the appropriate component | - | Pretty basic ICA evaluation paper; JADE algorithm was used; no useful comparison to other approaches that could be added to the list
Jung01 | 725 | Imaging Brain Dynamics Using Independent Component Analysis | Jung TP, Makeig S, McKeown MJ, Bell AJ, Lee TW, Sejnowski TJ. | Proc IEEE Inst Electr Electron Eng | cs | attention strategy | Outlining the assumptions underlying ICA in real EEG data. The subjects were exposed to visual stimuli and they had to press a button in 25% of the cases. | 29-channel EEG system (Electrocap, Inc.) 10-20 System, referenced to the right mastoid, 512Hz sample rate | 10 subjects (8 males) right-handed with normal or corrected vision | 10 | - |  | no | There are 4 assumptions underlying ICA, which can not assumed in practise: 1) Signal conduction times are equal and the summation linear. 2) Spatial projections of components are fixed in time and conditions. 3) Temporal independence of the components. 4) Statistical distributions of the component activation are not Gaussian. Generally, ICA is still very effective and stable for EEG and fMRI.
Total progress | 77 |  |  |  | 77
100%

## Pipelines
|  |  | Device | [5 empty] | Filters |  |  |  | Aggregation |  | Decomposition | [13 empty] | Detection | [5 empty] | Correction |  |  | Special algorithms | [7 empty] | Classification
Reference | Pipeline ID (unique per paper) | is new | EEG (#ch) | electrode type (w,g,d) | MEG | ECG | EMG | Others | Filter (HP, LP, BP, BR) | MA | AF | Other | Epoch width (s) | Feature Extraction (#) | Fourier Transf. | Wavelet Transf. | Welch | LSP | PCC | CCA | ICA | NMF | PCA | t-SNE | EMD | IVA | CSP | Other | Linear Regression | DA | SVM | Spatial sph. spline | GMM | Other | Autoencoder | GAN | Other | MARA | DSS | DFA | FASTER | ADJUST | HAPPE | ERASE | Other | DA | KNN | KMC | ED | NB | DT | RF | SVM | NN | GMM | MEC | Other | fit score (0-2) | online score (0-2) | details
Hossain22 | WPD(db1) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db1 | [41 empty] | 2 | 2 | Online score depends on concrete implementation of WPD but should be real-time compatible by nature. The authors claim novel approach, but questionable that they invented WPD.
Hossain22 | WPD(db2) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db2 | [41 empty] | 2 | 2
Hossain22 | WPD(db3) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db3 | [41 empty] | 2 | 2
Hossain22 | WPD(sym4) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: sym4 | [41 empty] | 2 | 2
Hossain22 | WPD(sym5) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: sym5 | [41 empty] | 2 | 2
Hossain22 | WPD(sym6) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: sym6 | [41 empty] | 2 | 2
Hossain22 | WPD(coif1) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: coif1 | [41 empty] | 2 | 2
Hossain22 | WPD(coif2) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: coif2 | [41 empty] | 2 | 2 | The authors claim novel approach, but questionable that they invented WPD
Hossain22 | WPD(coif3) | 1 | 2 |  |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: coif3 | [41 empty] | 2 | 2
Hossain22 | WPD(fk4) | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk4 | [41 empty] | 2 | 2
Hossain22 | WPD(fk6) | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk6 | [41 empty] | 2 | 2
Hossain22 | WPD(fk8) | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk8 | [41 empty] | 2 | 2
Hossain22 | WPD(db1)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db1 |  |  |  | 3 | [37 empty] | 2 | 2
Hossain22 | WPD(db2)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db2 |  |  |  | 3 | [37 empty] | 2 | 2
Hossain22 | WPD(db3)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: db3 |  |  |  | 3 | [37 empty] | 2 | 2
Hossain22 | WPD(fk4)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk4 |  |  |  | 3 | [37 empty] | 2 | 2
Hossain22 | WPD(fk6)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk6 |  |  |  | 3 | [37 empty] | 2 | 2
Hossain22 | WPD(fk8)-CCA | 1 | 2 | - |  |  |  | acc | 1: NOTCH (3rd order Butterworth; center-f at 50 Hz) + Baseline Drift Removal | [6 empty] | 2: fk8 |  |  |  | 3 | [37 empty] | 2 | 2
Abu22 | FastICA-SVM | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: Fast ICA | [31 empty] | 3 |  |  |  |  | 1 | 2 | implementations are fast enough for RT scenario but unclear if working well enough to be a good fit
Abu22 | FastICA-KNN | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: Fast ICA | [25 empty] | 3 | [10 empty] | 1 | 2
Abu22 | FastICA-DT | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: Fast ICA | [29 empty] | 3 | [6 empty] | 1 | 2
Abu22 | FastICA-DA | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: Fast ICA | [24 empty] | 3 | [11 empty] | 1 | 2
Abu22 | FastICA-NB | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: Fast ICA | [28 empty] | 3 | [7 empty] | 1 | 2
Abu22 | wICA-SVM | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: wICA | [31 empty] | 3 |  |  |  |  | 1 | 2
Abu22 | wICA-KNN | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: wICA | [25 empty] | 3 | [10 empty] | 1 | 2
Abu22 | wICA-DT | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: wICA | [29 empty] | 3 | [6 empty] | 1 | 2
Abu22 | wICA-DA | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: wICA | [24 empty] | 3 | [11 empty] | 1 | 2
Abu22 | wICA-NB | 0 | 64 | g |  |  |  |  | 1: NOTCH (center-f at 50 Hz) + BP 0.1 Hz - 40 Hz | [11 empty] | 2: wICA | [28 empty] | 3 | [7 empty] | 1 | 2
Flo22 | APICE | 1 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10), 5: HP 0.2 (epoching) |  |  |  | 2: 1s | [9 empty] | 3: target PCA | [8 empty] | 4 | [25 empty] | 0 | 1 | multiple cycles of artifact removal over all and epoched data
Flo22 | APICEa | 1 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10), 3: HP 0.2 (epoching) |  |  |  | 2: 1s | [9 empty] | 3: target PCA | [8 empty] | 4 | [25 empty] | 0 | 1 | Filtering on continuous data and not in epoched, therefore online score 1
Flo22 | APICE+wICA+iMARA | 1 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10), 9: HP 0.2 (epoching) |  |  |  | 2: 1s | [7 empty] | 5: wICA, 6 |  | 3: target PCA | [8 empty] | 4, 8, 10 | [5 empty] | 7: iMARA | [19 empty] | 0 | 1 | Overview in the Paper is very good, everything works sequential here, HP before epoching
Flo22 | APICE+DSS | 1 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10), 5: HP 0.2 (epoching) |  |  |  | 2: 1s | [9 empty] | 3: target PCA | [8 empty] | 4 | [6 empty] | 5 | [18 empty] | 0 | 1
Flo22 | Standard Pipeline (Winkler 2009) | 0 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10) |  |  |  | 2: 1s | [18 empty] | 3 | [25 empty] | 0 | 1
Flo22 | MADE (Debnath 2020) | 0 | 128 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10) |  |  |  | 2: 1s | [7 empty] | 4 | [10 empty] | 3 | [8 empty] | 3 | 5: adjusted-ADJUST | [15 empty] | 0 | 1
Kumaravel22 | NEAR (ASR R) | 1 | 125 | [5 empty] | 1: BP 0.3-30 (FIR) |  |  |  | 2: 10s | [10 empty] | 4: ASR | [9 empty] | 3: LOF | [23 empty] | 0 | 0
Kumaravel22 | NEAR (ASR C) | 1 | 125 | [5 empty] | 1: BP 0.3-30 (FIR) |  |  |  | 2: 10s | [10 empty] | 4: ASR | [9 empty] | 3: LOF | [23 empty] | 2 | 0
Kumaravel22 | CRD | 0 | 125 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10) |  |  |  | 2: 10s | [31 empty] | 3: clean_rawdata | [12 empty] | 1 | 0
Kumaravel22 | MADE | 0 | 125 | - |  |  |  |  | 1: BP 0.1-40 (FIR, bands: 0.1,10) |  |  |  | 2: 10s | [7 empty] | 4 | [10 empty] | 3 | [8 empty] | 3 | 5: adjusted-ADJUST | [15 empty] | 1 | 0 | MADE (Debnath 2020)
Kumaravel22 | HAPPE | 0 | 125 | [5 empty] | 1: BP 0.1-40 (FIR, bands: 0.1,10) |  |  |  | 2: 10s | [7 empty] | 3: wICA | [16 empty] | 4 | [19 empty] | 1 | 0
Soghoyan21 | LR | 0 | 31 | [5 empty] | 1: BP 0.1-40Hz |  |  |  | 2: 1.3s | 5 | [6 empty] | 4: FastICA, 4: AMICA | [7 empty] | 5: LR | [28 empty] | 1 | 0 | features: kurtosis, maximum epoch variance, spatial average difference (SAD), spatial eye distance (SED), Myogenic identification feature (MIF), correlations, mu and alpha topography, average magnitude in alpha band.
Soghoyan21 | SVM | 0 | 31 | [5 empty] | 1: BP 0.1-40Hz |  |  |  | 2: 1.3s | 5 | [6 empty] | 4: FastICA, 4: AMICA | [9 empty] | 5: SVM | [26 empty] | 1 | 0 | very good source of data
Soghoyan21 | XGB | 0 | 31 | [5 empty] | 1: BP 0.1-40Hz |  |  |  | 2: 1.3s | 5 | [6 empty] | 4: FastICA, 4: AMICA | [7 empty] | 5: XGB | [28 empty] | 1 | 0
Karpiel21 | Original EEG | 0 | 6 | [54 empty] | 0 | 2
Karpiel21 | FIR | 0 | 6 | [5 empty] | 1: BP 5-40Hz (FIR) | [48 empty] | 1 | 2 | only filters are compared
Karpiel21 | IIR | 0 | 6 | [5 empty] | 1: BP 5-40Hz (IIR) | [48 empty] | 1 | 2 | weak results
Karpiel21 | FFT | 0 | 6 | [5 empty] | 1: BP 5-40Hz | [48 empty] | 1 | 2
Karpiel21 | NOTCH | 0 | 6 | [5 empty] | 1: BR 46.5-127Hz (NOTCH) | [48 empty] | 1 | 2
Beach21 | Original EEG | 0 | 4 | g |  |  |  |  | 1: BP 0.16-70 Hz (Butterworth), 2: NOTCH 47.5-52.5 Hz | [48 empty] | 2 | 0
Beach21 | AF acc | 1 | 4 | g |  |  |  | accelerometer | 1: BP 0.16-70 Hz (Butterworth), 2: NOTCH 47.5-52.5 Hz |  |  | 4 | 3: integrate the signal | [44 empty] | 2 | 0
Beach21 | AF gyro | 1 | 4 | g |  |  |  | gyroscope | 1: BP 0.16-70 Hz (Butterworth), 2: NOTCH 47.5-52.5 Hz |  |  | 4 | 3: Savitzky-Golay filter | [44 empty] | 2 | 0
Dai21 | EEMD-CCA | 0 | 22 | [5 empty] | 1: BP 0.5-30Hz | [10 empty] | 3: CCA |  |  |  |  | 2: EEMD | [32 empty] | 2 | 0
Dai21 | sEEMD-CCA | 1 | 22 | [5 empty] | 1: BP 0.5-30Hz | [10 empty] | 3: CCA |  |  |  |  | 2: sEEMD | [32 empty] | 2 | 0
Dai21 | EEMD-ICA | 0 | 22 | [5 empty] | 1: BP 0.5-30Hz | [11 empty] | 3: ICA |  |  |  | 2: EEMD | [32 empty] | 2 | 0
Tost21 | distribution-based double-rejection | 0 | 20 |  |  |  |  | accelerometer | 1: BP 1.5-80Hz, 3: BP 1-10Hz |  |  |  | 2: 5s | 4: mean+std, 5: mean+std | [43 empty] | 0 | 0 | no real comparisions between pipelines
Tost21 | energy double-rejection | 0 | 20 |  |  |  |  | accelerometer | 1: BP 1.5-80Hz, 3: BP 1-10Hz |  |  |  | 2: 5s | 4: entropy, 5: entropy | [43 empty] | 0 | 0 | artifacts were only detected and not corrected
Tost21 | hybrid double-rejection | 1 | 20 |  |  |  |  | accelerometer | 1: BP 1.5-80Hz, 3: BP 1-10Hz |  |  |  | 2: 5s | 4: entropy, 5: mean+std | [43 empty] | 0 | 0
Kaur21 | DWT | 0 | - | - |  |  |  |  | 1: Filtering (no details provided) |  |  |  | 2: Epoching N sampled data (no details provided) |  |  | 3: DWT | [41 empty] | 1 | 1 | ignored the classification pipelines as they are used for depression detection only
Kaur21 | EMD | 0 | - | - |  |  |  |  | 1: Filtering (no details provided) |  |  |  | 2: Epoching N sampled data (no details provided) | [11 empty] | 3 | [32 empty] | 1 | 1
Kaur21 | EMD-DWT | 0 | - | - |  |  |  |  | 1: Filtering (no details provided) |  |  |  | 2: Epoching N sampled data (no details provided) |  |  | 4: DWT | [8 empty] | 3 | [32 empty] | 1 | 1
Kaur21 | EMD-DFA-WPD | 1 | - | - |  |  |  |  | 1: Filtering (no details provided) |  |  |  | 2: Epoching N sampled data (no details provided) |  |  | 5 | [8 empty] | 3 | [14 empty] | 4 | [17 empty] | 1 | 1
Rosanne21 | AF-RF: Original EEG | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [37 empty] | 4 | [5 empty] | 2 | 1 | only pipelines are compared here, for individual feature results see paper for more details
Rosanne21 | AF-RF: ASR | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [8 empty] | 4: ASR | [28 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | AF-RF: ADJUST | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [27 empty] | 4 | [9 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | AF-RF: HAPPE | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [28 empty] | 4 | [8 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | AF-RF: ASR_ADJUST | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [8 empty] | 4: ASR | [18 empty] | 5 | [9 empty] | 6 | [5 empty] | 2 | 1
Rosanne21 | AF-RF: ASR_wICA | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [6 empty] | 5: wICA |  | 4: ASR | [28 empty] | 6 | [5 empty] | 2 | 1
Rosanne21 | AF-RF: wICA | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [6 empty] | 4: wICA | [30 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | AF-SVM: Original EEG | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [38 empty] | 4 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: ASR | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [8 empty] | 4: ASR | [29 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: ADJUST | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [27 empty] | 4 | [10 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: HAPPE | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [28 empty] | 4 | [9 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: ASR_ADJUST | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [8 empty] | 4: ASR | [18 empty] | 5 | [10 empty] | 6 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: ASR_wICA | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [6 empty] | 5: wICA |  | 4: ASR | [29 empty] | 6 |  |  |  |  | 2 | 1
Rosanne21 | AF-SVM: wICA | 0 | 8 | d |  |  |  | accelerometer | 1: BP 1-45Hz (FIR) |  | 2: adaptive filter |  |  | 3: PSD, PMSC, AM, PMSC-AM | [6 empty] | 4: wICA | [31 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | RF: Original EEG | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [37 empty] | 3 | [5 empty] | 2 | 1
Rosanne21 | RF: ASR | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [8 empty] | 3: ASR | [28 empty] | 4 | [5 empty] | 2 | 1
Rosanne21 | RF: ADJUST | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [27 empty] | 3 | [9 empty] | 4 | [5 empty] | 2 | 1
Rosanne21 | RF: HAPPE | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [28 empty] | 3 | [8 empty] | 4 | [5 empty] | 2 | 1
Rosanne21 | RF: ASR_ADJUST | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [8 empty] | 3: ASR | [18 empty] | 4 | [9 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | RF: ASR_wICA | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [6 empty] | 4: wICA |  | 3: ASR | [28 empty] | 5 | [5 empty] | 2 | 1
Rosanne21 | RF: wICA | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [6 empty] | 3: wICA | [30 empty] | 4 | [5 empty] | 2 | 1
Rosanne21 | SVM: Original EEG | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [38 empty] | 3 |  |  |  |  | 2 | 1
Rosanne21 | SVM: ASR | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [8 empty] | 3: ASR | [29 empty] | 4 |  |  |  |  | 2 | 1
Rosanne21 | SVM: ADJUST | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [27 empty] | 3 | [10 empty] | 4 |  |  |  |  | 2 | 1
Rosanne21 | SVM: HAPPE | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [28 empty] | 3 | [9 empty] | 4 |  |  |  |  | 2 | 1
Rosanne21 | SVM: ASR_ADJUST | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [8 empty] | 3: ASR | [18 empty] | 4 | [10 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | SVM: ASR_wICA | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [6 empty] | 4: wICA |  | 3: ASR | [29 empty] | 5 |  |  |  |  | 2 | 1
Rosanne21 | SVM: wICA | 0 | 8 | d |  |  |  |  | 1: BP 1-45Hz (FIR) |  |  |  |  | 2: PSD, PMSC, AM, PMSC-AM | [6 empty] | 3: wICA | [31 empty] | 4 |  |  |  |  | 2 | 1
Nahmias21 | DCNN | 1 | 58 | - |  |  |  |  | 1 | [44 empty] | 2: DCNN |  |  |  | 2 | 1
Saba-Sadiya21 | HBOS | 1 | 32 | - | [8 empty] | 1: 0.2 | 2: 58 | [13 empty] | 3: HBOS | [6 empty] | 4 | [17 empty] | 5 |  |  |  |  | 2 | 1 | This paper does a artifact Correction and not a simple detection and removal of the difficult epoch
Saba-Sadiya21 | LOF | 1 | 32 | - | [8 empty] | 1: 0.2 | 2: 58 | [13 empty] | 3: LOF | [6 empty] | 4 | [17 empty] | 5 |  |  |  |  | 2 | 1
Saba-Sadiya21 | ABOD | 1 | 32 | - | [8 empty] | 1: 0.2 | 2: 58 | [13 empty] | 3: ABOD | [6 empty] | 4 | [17 empty] | 5 |  |  |  |  | 2 | 1
Saba-Sadiya21 | OCSVN | 1 | 32 | - | [8 empty] | 1: 0.2 | 2: 58 | [16 empty] | 3: OCSVM |  |  |  | 4 | [17 empty] | 5 |  |  |  |  | 2 | 1
Saba-Sadiya21 | LSCP (2*HBOS+OCSVM) | 1 | 32 | - | [8 empty] | 1: 0.2 | 2: 58 | [13 empty] | 3: LSCP (2*HBOS+OCSVM) |  |  | 3: OCSVM |  |  |  | 4 | [17 empty] | 5 |  |  |  |  | 2 | 1
Saba-Sadiya21 | AUTO | 1 | 32 | - | [8 empty] | 1: 0.2 | [21 empty] | 2 | [17 empty] | 3 |  |  |  |  | 2 | 1
Saba-Sadiya21 | PCA | 1 | 32 | - | [8 empty] | 1: 0.2 | [9 empty] | 2 | [29 empty] | 3 |  |  |  |  | 2 | 1
Saba-Sadiya21 | VAE | 1 | 32 | - | [8 empty] | 1: 0.2 | [21 empty] | 2: VAE | [17 empty] | 3 |  |  |  |  | 2 | 1
Saba-Sadiya21 | GAAL | 1 | 32 | - | [8 empty] | 1: 0.2 | [22 empty] | 2: GAAL | [16 empty] | 3 |  |  |  |  | 2 | 1
Saba-Sadiya21 | MGAAL | 1 | 32 | - | [8 empty] | 1: 0.2 | [22 empty] | 2: MGAAL | [16 empty] | 3 |  |  |  |  | 2 | 1
Saba-Sadiya21 | Original EEG | 0 | 32 | - | [8 empty] | 1: 0.2 | [39 empty] | 2 |  |  |  |  | 0 | 2
Li21 | Original EEG | 0 | 64 |  |  |  | 16 |  | 1: BP 3-100Hz | [48 empty] | 0 | 2
Li21 | ERASE-Real-EMG | 1 | 64 |  |  |  | 16 |  | 1: BP 3-100Hz | [24 empty] | 3: RMS |  |  | 4: max. abs. of coefficients | [6 empty] | 2 | [13 empty] | 2 | 1
Li21 | ERASE-Simulated-EMG | 1 | 64 |  |  |  | 16 |  | 1: BP 3-100Hz | [24 empty] | 3: RMS |  |  | 4: max. abs. of coefficients | [6 empty] | 2 | [13 empty] | 2 | 1
Li21 | Conventional-ICA | 0 | 64 |  |  |  | 16 |  | 1: BP 3-100Hz | [11 empty] | 2 | [12 empty] | 3: RMS |  |  | 4: max. abs. of coefficients | [20 empty] | 1 | 1
Anders20 | ASR | 0 | 64 | [5 empty] | 1: BP 3-30Hz | [11 empty] | 3: AMICA |  | 2: ASR | [10 empty] | 4: RELICA, 4: IClabel | [23 empty] | 2 | 1 | This pipeline was not tested against other pipelines
Zhang20 | PEDOT:PSS | 1 | 2 | 3D-PWS | [53 empty] | 0 | 2 | This paper proposes an electrode type, so no further pipeline description is needed
Leach20 | adjusted ADJUST | 1 | 64 | w |  |  |  |  | 1: BP 0.3-49Hz, 3: HP 1Hz |  |  |  | 4: 1s | 5: 2 | 5: power spectrum | [5 empty] | 6 | [10 empty] | 8 | [8 empty] | 2 | 7: adjusted-ADJUST | [15 empty] | 1 | 1 | Any epochs where the channel voltage exceeded +- 1mV or power within the 20-40Hz band exceeded -100dB or 30dB were deleted
Leach20 | original ADJUST | 0 | 64 | w |  |  |  |  | 1: BP 0.3-49Hz, 3: HP 1Hz |  |  |  | 4: 1s | 5: 2 | 5: power spectrum | [5 empty] | 6 | [10 empty] | 8 | [8 empty] | 2 | 7 | [15 empty] | 1 | 1
Leach20 | ICLabel | 0 | 64 | w |  |  |  |  | 1: BP 0.3-49Hz, 3: HP 1Hz |  |  |  | 4: 1s | 5: 2 | 5: power spectrum | [5 empty] | 6 | [10 empty] | 8 | [8 empty] | 2 |  |  |  | 7: IClabel | [12 empty] | 1 | 1
Leach20 | No correction | 0 | 64 | w |  |  |  |  | 1: BP 0.3-49Hz, 3: HP 1Hz |  |  |  | 4: 1s | 5: 2 | 5: power spectrum | [16 empty] | 7 | [8 empty] | 2 | 6: adjusted-ADJUST | [15 empty] | 1 | 2 | Any epochs where the channel voltage exceeded +- 1mV or power within the 20-40Hz band exceeded -100dB or 30dB were deleted
Saini20 | EEMD-CCA | 0 | - | - | [15 empty] | 1: CCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1
Saini20 | EEMD | 0 | - | - | [20 empty] | 1: EEMD | [32 empty] | 2 | 1
Saini20 | EMD-CCA | 0 | - | - | [15 empty] | 1: CCA |  |  |  |  | 1: EMD | [32 empty] | 2 | 1
Saini20 | Wavelet Denoising | 0 | - | - | [11 empty] | 1: DWT | [41 empty] | 2 | 1
Saini20 | EEMD-MCCA | 0 | - | - | [15 empty] | 1: MCCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1
Saini20 | VMD-cross | 1 | - | - | [23 empty] | 1: VMD | [5 empty] | 2: zero cross count threshold |  |  | 3: Retention/Rejection | [20 empty] | 2 | 1 | [1] Does not mention any pre-filtering of signals, only VMD center frequencies; [2] other pipelines (EEMD-CCA, EEMD, EMD-CCA, Wavelet Denoising, EEMD-MCCA) not listed here because not described in paper
Mucarquer20 | SuBAR | 0 | 32 | - |  |  | 0 |  | 1: BP 0.1 - 256 Hz |  |  |  | 2: 4 | [23 empty] | 3: SuBAR
Mucarquer20 | EEMD-CCA | 0 | 32 | - |  |  | 0 |  | 1: BP 0.1 - 256 Hz |  |  |  | 2: 4 | [6 empty] | 4: CCA |  |  |  |  | 3: EEMD
Mucarquer20 | EEMD-CCA-2EMG | 1 | 32 | - |  |  | 2 |  | 1: BP 0.1 - 256 Hz |  |  |  | 2: 4 | [6 empty] | 4: CCA |  |  |  |  | 3: EEMD
Mucarquer20 | EEMD-CCA-4EMG | 1 | 32 | - |  |  | 4 |  | 1: BP 0.1 - 256 Hz |  |  |  | 2: 4 | [6 empty] | 4: CCA |  |  |  |  | 3: EEMD
Mucarquer20 | EEMD-CCA-8EMG | 1 | 32 | - |  |  | 8 |  | 1: BP 0.1 - 256 Hz |  |  |  | 2: 4 | [6 empty] | 4: CCA |  |  |  |  | 3: EEMD
Nordin20 | Single-Layer-AMICA |  | 128 | g |  |  | 0 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  |  | 2: Channel rejection (Kurtosis and standard deviation) | [6 empty] | 3: AMICA | [26 empty] | 4 | [9 empty] | 2 | 2 | high value findings for walking
Nordin20 | Dual-Layer-AMICA |  | 128 | g |  |  | 8 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  |  | 2: Channel rejection (Kurtosis and standard deviation) | [6 empty] | 3: AMICA | [26 empty] | 4 | [9 empty] | 2 | 2
Nordin20 | Noise-Cancel-AMICA |  | 128 | g |  |  | 8 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  |  | 2: Channel rejection (Kurtosis and standard deviation) | [6 empty] | 3: AMICA | [26 empty] | 4 | [9 empty] | 2 | 2
Nordin20 | ASR-AMICA |  | 128 | g |  |  | 0 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  | 3: ASR | 2: Channel rejection (Kurtosis and standard deviation) | [6 empty] | 4: AMICA | [26 empty] | 5 | [9 empty] | 2 | 2
Nordin20 | ASR-merged-AMICA |  | 128 | g |  |  | 8 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  | 3: ASR | 2: Channel rejection (Kurtosis and standard deviation) | [6 empty] | 4: AMICA | [26 empty] | 5 | [9 empty] | 2 | 2
Nordin20 | PCA-CCA-merged-AMICA |  | 128 | g |  |  | 8 | additionally 40 noise electrodes, mechanically coupled | 1: HP 1Hz |  |  |  |  | 2: Channel rejection (Kurtosis and standard deviation) | [5 empty] | 4 | 5: AMICA |  | 3 | [24 empty] | 6 | [9 empty] | 2 | 2
Liu19 | CCA | 0 | 19 | [5 empty] | 1: HP 1Hz, 2: BP 0.3-35Hz |  |  |  |  | 3: 10s | [5 empty] | 4 | [37 empty] | 2 | 0 | The epochs are far to long
Liu19 | EEMD-CCA | 0 | 19 | [5 empty] | 1: HP 1Hz, 2: BP 0.3-35Hz |  |  |  |  | 3: 10s | [5 empty] | 5 |  |  |  |  | 4: EEMD | [32 empty] | 2 | 0
Liu19 | SSH-CCA | 1 | 19 | [5 empty] | 1: HP 1Hz, 2: BP 0.3-35Hz |  |  |  |  | 3: 10s | [5 empty] | 6 | [7 empty] | 4: SSA, 5: SSA | [29 empty] | 2 | 0 | SSA twice is enough, they optimize also the parameters
Benda19 | MC-Wiener-Filtered | 1 | 14 | g |  |  |  |  | 1: BP 2Hz -60 Hz, Notch (45Hz - 55Hz) |  |  | 2: Multi-Channel Wiener Filter |  |  | 3: FFT | [18 empty] | 4: Center of Gravity | [21 empty] | 5 |  | 0 | 2 | Has been shown to work in an online scenario
Benda19 | MC-Wiener-Unfiltered | 1 | 14 | g |  |  |  |  | 1: BP 2Hz -60 Hz, Notch (45Hz - 55Hz) | [5 empty] | 2: FFT | [18 empty] | 3: Center of Gravity | [21 empty] | 4 |  | 0 | 2
Val-Calvo19 | EAWICA-ALL | 0 | 8 | [5 empty] | 1: BP 1-4, BP 4-8, BP 8-16, BP 16-30, BP 30-50 |  |  |  | 2: 6s | 3: 200 |  | 4 |  |  |  |  | 4: EAWICA | [36 empty] | 1 | 2 | Features: Differential Entropy, Amplitude Envelope, Perosian Fractal Dimension, Higuchi Fractal Dimenstion, Fisher Information
Val-Calvo19 | ICA-W-ALL | 0 | 8 | [5 empty] | 1: BP 1-4, BP 4-8, BP 8-16, BP 16-30, BP 30-50 |  |  |  | 2: 6s | 3: 200 |  | 4 |  |  |  |  | 4: ICA-W | [36 empty] | 1 | 2 | Is not evaluated explicitely on motion artifacts
Val-Calvo19 | EAWICA-DELTA | 0 | 8 | [5 empty] | 1: BP 1-4 |  |  |  | 2: 6s | 3: 40 |  | 4 |  |  |  |  | 4: EAWICA | [36 empty] | 1 | 2
Val-Calvo19 | ICA-W-DELTA | 0 | 8 | [5 empty] | 1: BP 1-4 |  |  |  | 2: 6s | 3: 40 |  | 4 |  |  |  |  | 4: ICA-W | [36 empty] | 1 | 2
Val-Calvo19 | Linear dynamic system | 0 | 8 | [5 empty] | 1: BP 1-4, BP 4-8, BP 8-16, BP 16-30, BP 30-50 |  |  | 5: Linear dynamic system | 2: 6s | 3: 200 |  | 4 |  |  |  |  | 4: EAWICA | [36 empty] | 1 | 2
Val-Calvo19 | Savetzky-Golay | 0 | 8 | [5 empty] | 1: BP 1-4, BP 4-8, BP 8-16, BP 16-30, BP 30-50 |  |  | 5: Savetzky-Golay filter | 2: 6s | 3: 200 |  | 4 |  |  |  |  | 4: EAWICA | [36 empty] | 1 | 2
Val-Calvo19 | Proposed Method | 1 | 8 | [5 empty] | 1: BP 1-4, BP 4-8, BP 8-16, BP 16-30, BP 30-50 |  |  | 5: Savetzky-Golay filter | 2: 6s | 3: 200 |  | 4 |  |  |  |  | 4: EAWICA | [24 empty] | 6: QDA | 6 |  |  | 6 | 6 | 6 | 6: linear and radial |  |  |  | 6: Adaboost | 1 | 2 | The best classifactor was applied for inference
Saavedra19 | xDAWN | 0 | 10 | - |  |  |  |  | 1: Chebyshev bandpass filter (8th order; 0.1–60 Hz) + notch filter (60 Hz) | [27 empty] | 2: xDAWN | [20 empty] | 0 | 1
Saavedra19 | METS_SEWS_1 | 1 | 10 | - |  |  |  |  | 1: Chebyshev bandpass filter (8th order; 0.1–60 Hz) + notch filter (60 Hz) | [27 empty] | 2: METS | [7 empty] | 3: SEWS-1 | [12 empty] | 0 | 1
Saavedra19 | METS_SEWS_2 | 1 | 10 | - |  |  |  |  | 1: Chebyshev bandpass filter (8th order; 0.1–60 Hz) + notch filter (60 Hz) | [27 empty] | 2: METS | [7 empty] | 3: SEWS-2 | [12 empty] | 0 | 1
Mur19 | UAR | 1 | 64 | [9 empty] | 1: 0.6s (sliding window) | [7 empty] | 3: FastICA | [12 empty] | 2: UMED, 4: F-Test | [23 empty] | 0 | 2 | Proposed pipeline, assumes an artifact-free time before and after each corrected artifact and no bad channel.
Mur19 | MARA | 0 | 64 | [9 empty] | 1: 0.6s (sliding window) | [7 empty] | 2 | [16 empty] | 2 | [19 empty] | 1 | 1
Mur19 | SASICA | 0 | 64 | [9 empty] | 1: 0.6s (sliding window) | 2: 3 | [6 empty] | 3 | [20 empty] | 4 | [15 empty] | 1 | 1 | Features: Autocorrelation, Focal Components, Correlation between channels
Mur19 | wICA | 0 | 64 | [9 empty] | 1: 0.6s (sliding window) |  |  | 2: coif5 |  |  |  |  | 2: wICA | [36 empty] | 1 | 1
Blum19 | Original EEG | 0 | 24 | [5 empty] | 1: LP 40Hz, 2: HP 0.25Hz | [48 empty] | 0 | 2
Blum19 | ASR | 0 | 24 | [5 empty] | 1: LP 40Hz, 2: HP 0.25Hz | [13 empty] | 3: ASR | [34 empty] | 1 | 1
Blum19 | rASR | 1 | 24 | [5 empty] | 1: LP 40Hz, 2: HP 0.25Hz | [13 empty] | 3: rASR | [34 empty] | 1 | 1
Nordin19 | Infomax | 0 | 128 | g |  |  | 8 | additionally 40 noise electrodes (mechanically coupled) + camera + IMU + treadmill | 1: HP 1Hz |  |  |  | 4: 0.5s (sliding window with 94% overlap) |  | 5 | [5 empty] | 2: InfomaxICA | [6 empty] | 3: Cross Correlation to ground truth input measurement | [8 empty] | 6: artifact component rejection by flat power spectra (linear slope ≥ −0.06) | [7 empty] | 7: DIPFIT and Fieldtrip functions (EEGLAB) |  |  | 8 | [9 empty] | 2 | 2 | Pipeline in detail hard to enter here, after KMC for evaluation signal was low pass filtered (6 Hz cutoff, 4th order Butterworth)
Grosselin19 | LDA | 0 | 32 | w/d/w |  |  |  |  | 2: Notch (center-f at 50Hz) |  |  |  | 1: 1 (non-overlapping) |  | 3 | [12 empty] | 5: Spectral Distance |  | 4 | [27 empty] | 2 | 2 | Methods validated in real-time context
Grosselin19 | Linear SVM | 0 | 32 | w/d/w |  |  |  |  | 2: Notch (center-f at 50Hz) |  |  |  | 1: 1 (non-overlapping) |  | 3 | [12 empty] | 5: Spectral Distance | [24 empty] | 4 |  |  |  |  | 2 | 2
Grosselin19 | Euclidean | 0 | 32 | w/d/w |  |  |  |  | 2: Notch (center-f at 50Hz) |  |  |  | 1: 1 (non-overlapping) |  | 3 | [12 empty] | 5: Spectral Distance | [20 empty] | 4 | [8 empty] | 2 | 2
Grosselin19 | Weighted kNNs | 0 | 32 | w/d/w |  |  |  |  | 2: Notch (center-f at 50Hz) |  |  |  | 1: 1 (non-overlapping) |  | 3 | [12 empty] | 5: Spectral Distance | [18 empty] | 4 | [10 empty] | 2 | 2
Li18 | MEMD+SampEn | 1 | 8 | [5 empty] | 1: BP 3-30Hz (Butterworth) |  |  |  |  | 3: SampleEntropy, 4: energy and variance of WPD |  | 4: db-3 | [8 empty] | 2: NA-MEMD | [28 empty] | 5: MLP |  |  |  | 2 | 2 | Latency of 0.2s, facial expressions may induce EEG signals, maybe not the correct way
Chu18 | LSP-DBN | 1 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  |  |  |  | 3 | [35 empty] | 4: restricted Boltzmann machine (RBM) |  |  |  | 2 | 2
Chu18 | FFT-DBN | 0 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  | 3: FFT | [38 empty] | 4: restricted Boltzmann machine (RBM) |  |  |  | 2 | 2
Chu18 | Welch-DBN | 0 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  |  |  | 3 | [36 empty] | 4: restricted Boltzmann machine (RBM) |  |  |  | 2 | 2
Chu18 | LSP-SVM | 0 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  |  |  |  | 3 | [34 empty] | 4 |  |  |  |  | 2 | 2
Chu18 | FFT-SVM | 0 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  | 3: FFT | [37 empty] | 4 |  |  |  |  | 2 | 2
Chu18 | Welch-SVM | 0 | 64 | g |  |  |  |  | 1: LP 200Hz (cutoff-f=200-300Hz) + Notch (50Hz) + BP 8-35Hz (5th-order Butterworth; gain=1.5) |  |  |  | 2: 1 (80 % overlap, 0.2s step size) |  |  |  | 3 | [35 empty] | 4 |  |  |  |  | 2 | 2
San-Martin18 | Original EEG | 0 | 11 | d |  |  | 2 |  | 1: HP 0.25Hz, 2: BR 60Hz, 4: LP 40Hz |  |  |  | 3: 2s | [44 empty] | 0 | 2 | Pipelines were not tested explicitely
San-Martin18 | SASICA | 0 | 11 | d |  |  | 2 |  | 1: HP 0.25Hz, 2: BR 60Hz, 4: LP 40Hz |  | 7 |  | 3: 2s | 6 | [6 empty] | 5: InfomaxICA | [19 empty] | 7 | [16 empty] | 2 | 2 | Features: Autocorrelation, Focal components, focal trial activity, SNR
Sebek18 | Proposed method | 1 | 111 | - |  |  |  |  | 1: Notch (50 Hz) + HP 1 Hz | [11 empty] | 3: FastICA |  | 2 | [21 empty] | 4: Subspace Separation (new method) | 5: PSD Comparison | [11 empty] | 2 | 2 | Other piplines for comparison are not explained in detail within the paper therefore ignored
Arad18 | AMICA | 1 | 32 | w |  |  |  | 3-axis accelerometer | 2: HP 1 Hz + NOTCH (50 Hz) |  |  | 1: Channel Rejection with PREP |  |  | 5: FFT | [5 empty] | 3: AMICA | [15 empty] | 4: Removing via visual inspection | [7 empty] | 6: ASF (average stepping frequency) | [12 empty] | 2 | 0 | Channel rejection done via visual inspection, no full automatic pipeline
Arad18 | ASR | 0 | 32 | w |  |  |  | 3-axis accelerometer | 2: HP 1 Hz + NOTCH (50 Hz) |  |  | 1: Channel Rejection with PREP |  |  | 5: FFT | [7 empty] | 3: ASR | [13 empty] | 4: Removing via visual inspection | [7 empty] | 6: ASF (average stepping frequency) | [12 empty] | 2 | 0
Symeonidou18 | - | - | [55 empty] | - | - | Info: No pipeline is needed for this paper, hardware electrode comparison done
Stone18 | Fingerprint Method | 1 | 128 | g |  |  |  |  | 1: BP 0.3 - 100Hz, NOTCH (50Hz) | [13 empty] | 3: PCA for pre-whitening | [13 empty] | 2: Removing via visual inspection | [7 empty] | 4: Fingerprint Method | [7 empty] | 5: nonlinear binary SVMs |  |  |  |  | 2 | 1 | Only Fingerprint described, not compared to others
Frolich18 | Original EEG | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [48 empty] | 0 | 2
Frolich18 | TDSEP | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [11 empty] | 2: TDSEP | [12 empty] | 3: IC_MARC | [23 empty] | 2 | 1
Frolich18 | Infomax | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [11 empty] | 2: InfomaxICA | [12 empty] | 3: IC_MARC | [23 empty] | 2 | 1
Frolich18 | FastICA | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [11 empty] | 2: FastICA | [12 empty] | 3: IC_MARC | [23 empty] | 2 | 1
Frolich18 | HP-TDSEP | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth), 2: HP 14Hz | [11 empty] | 3: TDSEP | [12 empty] | 4: IC_MARC | [23 empty] | 2 | 1
Frolich18 | HP-Infomax | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth), 2: HP 14Hz | [11 empty] | 3: InfomaxICA | [12 empty] | 4: IC_MARC | [23 empty] | 2 | 1
Frolich18 | HP-FastICA | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth), 2: HP 14Hz | [11 empty] | 3: FastICA | [12 empty] | 4: IC_MARC | [23 empty] | 2 | 1
Frolich18 | SSD | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [18 empty] | 2: SSD | [29 empty] | 2 | 1
Frolich18 | Fourier-ICA | 0 | 64 | [5 empty] | 1: BP 2-45Hz (Butterworth) | [11 empty] | 2: Fourier-ICA | [12 empty] | 3: IC_MARC | [23 empty] | 2 | 1 | needs STFFT which makes it slower
Gabard-Durnam18 | HAPPE | 1 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 3: wICA + extended Infomax | [15 empty] | 2: Removing via User Selection | 4: MARA IC Rejection |  |  | 5: FASTER (Optional Segementation) |  |  |  | 6: spherical interpolation (Legendre polynomials up to the 7th order) | [12 empty] | 1 | 1 | complex pipelines, unfortunately difficult to enter here
Gabard-Durnam18 | ICA-MARA | 0 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 3: extended Infomax | [15 empty] | 2: Removing via User Selection | 4: MARA IC Rejection | [19 empty] | 1 | 1
Gabard-Durnam18 | Manual | 0 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 4: extended Infomax | [6 empty] | 3: Manual Channel Rejection | [8 empty] | 2: Removing via User Selection | 5: MARA IC Rejection | [19 empty] | 1 | 1
Gabard-Durnam18 | ASR | 0 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 4: extended Infomax |  | 3: ASR | [13 empty] | 2: Removing via User Selection | 5: MARA IC Rejection | [19 empty] | 1 | 1
Gabard-Durnam18 | ADJUST | 0 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 3: wICA + extended Infomax | [15 empty] | 2: Removing via User Selection |  |  |  |  | 4 | [15 empty] | 1 | 1
Gabard-Durnam18 | SASICA | 0 | 39 | [5 empty] | 1: HP 1 Hz+ LP 250 Hz + NOTCH (50/60 Hz) | [11 empty] | 3: wICA + extended Infomax | [15 empty] | 2: Removing via User Selection | [7 empty] | 4: SASICA | [12 empty] | 1 | 1
Gabard-Durnam18 | FASTER | 0 | 39 | [5 empty] | 2: HP 1 Hz + NOTCH (50/60 Hz) | [27 empty] | 1: Removing via User Selection |  |  |  | 3 | [16 empty] | 1 | 1
Gabard-Durnam18 | FASTER-MARA | 0 | 39 | [5 empty] | 2: HP 1 Hz + NOTCH (50/60 Hz) | [27 empty] | 1: Removing via User Selection |  |  |  | 3 | [16 empty] | 1 | 1
Tamburro18 | Gel, 20 IC | 1 | 128 | g |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (20) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Gel, 50 IC | 1 | 128 | g |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (50) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Gel, 80 IC | 1 | 128 | g |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (80) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Dry, 20 IC | 1 | 97 | d |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (20) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Dry, 50 IC | 1 | 97 | d |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (50) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Dry, 80 IC | 1 | 97 | d |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) |  |  |  | 4: 5s (sliding 1s overlap) or 2s | 5: 14 | [6 empty] | 3: InfomaxICA |  | 2: PCA (80) | [7 empty] | 6: binary SVM RBF kernel | [26 empty] | 2 | 1 | Used features: Kurtosis, Maximum epoch variance, Spatial average difference (SAD), Spacial eye difference (SED), PSD for 5 bands, cardiac identification feature (CIF), myogenic identification feature (MIF), eye movement template correlation (CORR), eyeblink template correlation (CORR), entropy feature (EF)
Tamburro18 | Gel, Original EEG | 0 | 128 | g |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) | [48 empty] | 0 | 2
Tamburro18 | Dry, Original EEG | 0 | 97 | d |  |  |  |  | 1: BP 0.3-100, BR 49-51 (FIR) | [48 empty] | 0 | 2
Lin18 | CCA+GMM | 1 | 60 | [5 empty] | 1: BP 0.1-60 Hz |  |  |  | 2 | 3: 10 (FFT) | [5 empty] | 2 | [12 empty] | 4 | [24 empty] | 1 | 1 | Features like CCA-components, fractal dimension, kurtosis etc. are well suited. FFT within feature extraction
Lin18 | ASR | 0 | 60 | [5 empty] | 1: BP 0.1-60 Hz |  |  |  | 2 | 3: 10 (FFT) | [8 empty] | 2: ASR | [9 empty] | 4 | [24 empty] | 1 | 1
Delisle-Rodriguez17 | Filter Ad WAR | 1 | 64 | [5 empty] | 1: BP 0.15-200Hz, NOTCH 50 |  | 4: Adaptive Spatial Filter |  | 2: 6s | 3: 12 |  |  |  |  | 4: CCC | [8 empty] | 5: Weighted Average Reference | [24 empty] | 6 |  |  |  |  | 1 | 0 | Features like fractal dimensions, reference-free EEG, and power density spectra bands from 0.1-70Hz
Delisle-Rodriguez17 | Filter WAR | 0 | 64 | [5 empty] | 1: BP 0.15-200Hz, NOTCH 51 |  |  |  | 2: 6s | 3: 12 | [13 empty] | 4: Weighted Average Reference | [24 empty] | 5 |  |  |  |  | 1 | 0
Delisle-Rodriguez17 | Ad LAR | 1 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 52 |  | 4: Adaptive Spatial Filter |  | 2: 0.25s | 3: 12 |  |  |  |  | 4: CCC | [8 empty] | 5: Local Average Reference | [24 empty] | 6 |  |  |  |  | 1 | 1 | Not directly shown to correct motion artifacts
Delisle-Rodriguez17 | Ad WAR | 1 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 53 |  | 4: Adaptive Spatial Filter |  | 2: 0.25s | 3: 12 |  |  |  |  | 4: CCC | [8 empty] | 5: Weighted Average Reference | [24 empty] | 6 |  |  |  |  | 1 | 1
Delisle-Rodriguez17 | LAR | 0 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 54 |  |  |  | 2: 0.25s | 3: 12 | [13 empty] | 4: Local Average Reference | [24 empty] | 5 |  |  |  |  | 1 | 1
Delisle-Rodriguez17 | WAR | 0 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 55 |  |  |  | 2: 0.25s | 3: 12 | [13 empty] | 4: Weighted Average Reference | [24 empty] | 5 |  |  |  |  | 1 | 1
Delisle-Rodriguez17 | Riemann Kernel | 0 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 56 |  |  |  | 2: 0.25s | 3: 12 | [13 empty] | 4: Riemann Kernel | [24 empty] | 5 |  |  |  |  | 1 | 1
Delisle-Rodriguez17 | CSP | 0 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 57 |  |  |  | 2: 0.25s | 3: 12 | [12 empty] | 4: CSP | [25 empty] | 5 |  |  |  |  | 1 | 1
Delisle-Rodriguez17 | Filter-Bank CSP | 0 | 20 | g |  |  |  |  | 1: BP 0.15-200Hz, NOTCH 58 |  |  |  | 2: 0.25s | 3: 12 | [12 empty] | 4: Filter-Bank CSP | 4: Filter-Bank CSP | [24 empty] | 5 |  |  |  |  | 1 | 1
Wang17 |  | 1 | 3 | d (new approach) |  |  |  |  | 1: BP 1-40 Hz (6th order Butterworth) | [48 empty] | 0 | 2 | online suitable but no motion artifact removal but for the dry strechable concentric electrodes
Jas17 | Autoreject | 1 | 60 |  | 248 |  |  |  | 1: LP 40Hz, BP 0.5-60Hz | [24 empty] | 2: Autoreject | [23 empty] | 1 | 1 | as far as we could tell this does mainly detect artifacts rather than rejecting them
Jas17 | FASTER | 0 | 60 |  | 248 |  |  |  | 1: LP 40Hz, BP 0.5-60Hz | [31 empty] | 2 | [16 empty] | 1 | 1
Jas17 | SNS | 0 | 60 |  | 248 |  |  |  | 1: LP 40Hz, BP 0.5-60Hz | [18 empty] | 2: SNS | [29 empty] | 1 | 1
Jas17 | RANSAC | 0 | 60 |  | 248 |  |  |  | 1: LP 40Hz, BP 0.5-60Hz | [24 empty] | 2: RANSAC | [23 empty] | 1 | 1
Jas17 | NoReject | 0 | 60 |  | 248 |  |  |  | 1: LP 40Hz, BP 0.5-60Hz | [48 empty] | 0 | 2
Oliveira17 | CR1 | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz |  |  |  | 2: one step | 3: 4 | [6 empty] | 4 | [36 empty] | 2 | 0 | Features for channel removal: Range, Kurtosis, Correlation with other channels, Standard deviation
Oliveira17 | TCR | 1 | 256 | g |  |  |  |  | 1: HP 1 Hz | 4 |  |  | 2: one step, 5: one step | 3: 4 |  |  |  |  | 6 |  | 8 | [6 empty] | 7: ARNG | [29 empty] | 2 | 0 | Assumes that motion artifacts are very repetitive and form a "typical" motion artifact template. This is not the case in real life
DelPozo-Banos17 | None-WithLCF | 1 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [15 empty] | 3: LCF | [19 empty] | 4: spherical spline technique | 2 | 1
DelPozo-Banos17 | ADJUST-WithLCF | 1 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [15 empty] | 4: LCF |  |  |  |  | 3 | [14 empty] | 5: spherical spline technique | 2 | 1
DelPozo-Banos17 | FASTER-WithLCF | 1 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [15 empty] | 4: LCF |  |  |  | 3 | [15 empty] | 5: spherical spline technique | 2 | 1
DelPozo-Banos17 | ICAW-WithLCF | 1 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [6 empty] | 3: ICAW |  |  |  |  | 2: InfomaxICA | [15 empty] | 4: LCF | [19 empty] | 5: spherical spline technique | 2 | 1
DelPozo-Banos17 | None-WithoutLCF | 0 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [35 empty] | 3: spherical spline technique | 2 | 1
DelPozo-Banos17 | ADJUST-WithoutLCF | 0 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [20 empty] | 3 | [14 empty] | 4: spherical spline technique | 2 | 1
DelPozo-Banos17 | FASTER-WithoutLCF | 0 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [11 empty] | 2: InfomaxICA | [19 empty] | 3 | [15 empty] | 4: spherical spline technique | 2 | 1
DelPozo-Banos17 | ICAW-WithoutLCF | 0 | 128 | - |  |  |  | 3 * EOG channels | 1: HP 0.5Hz (FIR filter of order 99) + Hamming Window | [6 empty] | 3: ICAW |  |  |  |  | 2: InfomaxICA | [35 empty] | 4: spherical spline technique | 2 | 1
Weiss17 | AR2-CUDAICA | 1 | - | - |  |  |  |  | 1: BP 16-70Hz (FIR; 500th order) | 3: Moving Average |  |  |  |  | 2: Hilbert | [5 empty] | 4: CUDAICA | [35 empty] | 5: visual inspection | 1 | 1 | Final low pass filtering not shown in pipeline here, seizure onset classification was done through visual inspection of experts
Roy17 | EEMD-CCA-DWT | 0 | 16 | [12 empty] | 3: DWT |  |  |  | 2: CCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1
Roy17 | EEMD-GECCA-DWT | 1 | 16 | [12 empty] | 3: DWT |  |  |  | 2: GECCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1 | GECCA is 18% faster, the accuracy should also be better, but the ROC shows that this depends on the sensitivity
Roy17 | EEMD-CCA-SWT | 0 | 16 | [12 empty] | 3: SWT |  |  |  | 2: CCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1
Roy17 | EEMD-GECCA-SWT | 1 | 16 | [12 empty] | 3: SWT |  |  |  | 2: GECCA |  |  |  |  | 1: EEMD | [32 empty] | 2 | 1
Melman16 | Robust Method w/Bayes | 1 | 128 | [5 empty] | 1: 0.01-100Hz, 2: NOTCH 60Hz |  |  |  | 3: 3s |  | 4: Median Power Density Spectrum by Bayes modelling | [42 empty] | 0 | 0
Melman16 | Standard Method w/Jackknife | 0 | 128 | [5 empty] | 1: 0.01-100Hz, 2: NOTCH 60Hz |  |  |  | 3: 3s |  | 4: Mean Power Density Spectrum | [42 empty] | 0 | 0
Oliveira16 | BioSemi | 0 | 248 | g |  |  |  |  | 1: BP 1-100Hz, 2: NOTCH |  |  |  | 4: 1.1s | [7 empty] | 3 | [36 empty] | 0 | 1
Oliveira16 | Cognionics-wet | 0 | 64 | g |  |  |  |  | 1: BP 1-131Hz, 2: NOTCH |  |  |  | 4: 1.1s | [7 empty] | 3 | [36 empty] | 0 | 1
Oliveira16 | Cognionics-dry | 0 | 64 | d |  |  |  |  | 1: BP 1-131Hz, 2: NOTCH |  |  |  | 4: 1.1s | [7 empty] | 3 | [36 empty] | 0 | 1
Cheveigne16 | STAR | 1 | 10 | [5 empty] | 1: HP 20 Hz |  |  |  | - | [14 empty] | 2: STAR | [29 empty] | 0 | 2 | Noted to be online suitable, but not if muscle artifacts exist
Cheveigne16 | JD (Cheveigne 2014) | 0 | 10 | [5 empty] | 1: HP 20 Hz |  |  |  | - | [14 empty] | 2: JD | [29 empty] | 0 | 1 | yes they really applied a HP at 20 Hz, as they focuse on the low gamma band
Cheveigne16 | STAR+JD | 1 | 10 | [5 empty] | 1: HP 20 Hz |  |  |  | - | [14 empty] | 2 STAR, 3: JD | [29 empty] | 0 | 1
Cheveigne16 | Original EEG | 0 | 10 | [5 empty] | 1: HP 20 Hz |  |  |  | - | [44 empty] | 0 | 2 | online score is 2, because it does nothing
Costa16 | Proposed Pipeline | 0 | 32 | [5 empty] | 1: BP 0.5-100Hz, 2: NOTCH 50Hz |  |  |  | 3: 0.5 s | 4: 1 | [43 empty] | 0 | 2 | Not sophisticated, uses only standard deviation as feature
Nathan16 | FFT Pipeline | 1 | 64 |  |  |  |  | accereration | 1 |  |  |  | 1: 0.5s |  | 2: FFT Spectrogram | [16 empty] | 3: Spline Interpolation | [25 empty] | 0 | 1 | They only analyse properties of the EEG during walking
Nathan16 | Wavelet Coherence | 1 | 64 |  |  |  |  | accereration | 1 |  |  |  | 1: 0.5s |  |  | 2: Wavelet Coherence | [15 empty] | 3: Spline Interpolation | [25 empty] | 0 | 1
Nathan16 | Wavelet Coherence + ASR | 1 | 64 |  |  |  |  | accereration | 1 |  |  |  | 1: 0.5s |  |  | 3: Wavelet Coherence | [6 empty] | 2: ASR | [8 empty] | 4: Spline Interpolation | [25 empty] | 0 | 1
Zou16 | Our Method | 1 | 8 | w |  |  |  | impedance measurement | 1: BP 0.5-50Hz |  |  |  | 2: 0.8s | 4: 8 | [6 empty] | 3: FastICA | [12 empty] | 5: impedance artifacts, 6: hierarchical clustering | [23 empty] | 1 | 1 | Features: Kurtosis, spatial, power bands, similarity over epochs. Interesing approach with the impedance measurement, but not verified
Zou16 | FASTER | 0 | 8 | w |  |  |  |  | 1: BP 0.5-50Hz |  |  |  | 2: 0.8s | [7 empty] | 3: FastICA | [19 empty] | 4 | [16 empty] | 2 | 1
Zou16 | ADJUST | 0 | 8 | w |  |  |  |  | 1: BP 0.5-50Hz |  |  |  | 2: 0.8s | [7 empty] | 3: FastICA | [20 empty] | 4 | [15 empty] | 2 | 1
Zou16 | Auto-mutual information | 0 | 8 | w |  |  |  |  | 1: BP 0.5-50Hz |  |  |  | 2: 0.8s | [7 empty] | 3: FastICA | [12 empty] | 4: hierarchical clustering | [23 empty] | 1 | 1
Zou16 | K-means with similarity | 0 | 8 | w |  |  |  |  | 1: BP 0.5-50Hz |  |  |  | 2: 0.8s | [7 empty] | 3: FastICA | [12 empty] | 4: k-means | [23 empty] | 1 | 1
Gordon15 | IMICA | 1 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [11 empty] | 2: IMICA | [36 empty] | 2 | 1
Gordon15 | Infomax | 0 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [11 empty] | 2: Infomax | [36 empty] | 2 | 1
Gordon15 | FastICA | 0 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [11 empty] | 2: FastICA | [36 empty] | 2 | 1
Gordon15 | JADE | 0 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [11 empty] | 2: JADE | [36 empty] | 2 | 1
Gordon15 | SOBI | 0 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [18 empty] | 2: SOBI | [29 empty] | 2 | 1
Gordon15 | PCA | 0 | 64 | g |  |  |  | EOG | 1: BP 1-50Hz | [13 empty] | 2 | [34 empty] | 2 | 1
Snyder15 | AMICA | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz |  |  |  | 2 | 3: power frequency spectra, average time course for a stride, and event-related spectral perturbations (ERSP) | [6 empty] | 4: AMICA | [36 empty] | 1 | 0 | Interesting to understand non-neuronal motion artifacts, non-usable for RT prototype
Snyder15 | Original EEG | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz | [48 empty] | 0 | 1
Kline15 | Moving Averaging | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz | 2: Moving Average | [47 empty] | 0 | 1
Kline15 | Wavelet | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz | [6 empty] | 2: Wavelet filter | [41 empty] | 0 | 1
Kline15 | Wavelet + Moving Averaging | 0 | 256 | g |  |  |  |  | 1: HP 1 Hz | 3: Moving Average | [5 empty] | 2: Wavelet filter | [41 empty] | 0 | 02
Raduntz15 | Proposed method | 1 | 25 | - |  |  |  |  | 1: Hamming Window + BP 0.5 Hz - 40 Hz (order: 100) |  |  |  |  | 3: LBP + gradient images + range filter + laplacian | [6 empty] | 2: Infomax | [24 empty] | 4: LDA | [11 empty] | 1 | 2
Mariani15 | multiscale entropy | 1 | 2 | [9 empty] | 1: 2s | 2: 1 | [43 empty] | 1 | 2 | Features: Multiscale Entropy
Zhang15 | WICA | 1 | 16 | [5 empty] | 1: BP 1-50Hz, 2: NOTCH 50Hz |  |  |  | 3: 5s | [7 empty] | 4: w-ICA | [31 empty] | 5: SVM (RBF) |  |  |  |  | 2 | 0 | Latency of 5s for 10s Windows
Zhang15 | Original EEG | 0 | 16 | [5 empty] | 1: BP 1-50Hz, 2: NOTCH 50Hz |  |  |  | 3: 5s | [39 empty] | 4: SVM (RBF) |  |  |  |  | 0 | 2
Bulea14 | Raw | 0 | 64 | [5 empty] | 1: HP 0.05Hz, 2: BP 0.1-4Hz | [20 empty] | 3: LFDA | [24 empty] | 4 |  |  | 0 | 2
Bulea14 | ASR | 0 | 64 | [5 empty] | 1: HP 0.05Hz, 3: BP 0.1-4Hz | [13 empty] | 2: ASR | [6 empty] | 4: LFDA | [24 empty] | 5 |  |  | 0 | 2 | ASR is not really tested to remove or correct myogenic artifacts
Chen14 | Original EEG | 0 | 21 | - |  |  |  |  | 1: BP 1-70 Hz
Chen14 | SCICA | 0 | 21 | - |  |  |  |  | 1: BP 1-70 Hz | [11 empty] | 2: SCICA | [36 empty] | 1 | 1 | Only the real data pipeline is shown here, simulated pipeline discarded
Chen14 | EEMD-ICA | 0 | 21 | - |  |  |  |  | 1: BP 1-70 Hz | [11 empty] | 3 |  |  |  | 2: EEMD | [32 empty] | 1 | 1
Chen14 | EEMD-CCA | 0 | 21 | - |  |  |  |  | 1: BP 1-70 Hz | [10 empty] | 3 |  |  |  |  | 2: EEMD | [32 empty] | 1 | 1
Chen14 | EEMD-IVA | 0 | 21 | - |  |  |  |  | 1: BP 1-70 Hz | [15 empty] | 2: EEMD | 3 | [31 empty] | 1 | 1
Chen14 | EEMD-JBSS | 1 | 21 | - |  |  |  |  | 1: BP 1-70 Hz | [15 empty] | 2: EEMD |  |  | 3: JBSS | [29 empty] | 1 | 1
Yilmaz14 | - | - | [57 empty] | Analysis done in Software Spike2, no need to add it here
Reis14 | -
Boudet14 | DAFOP | 1 | 19 | - | [8 empty] | 1 | [13 empty] | 2 | 3: frequency window decomp. | 4 | [28 empty] | 1 | 1
Boudet14 | BSS-CCA | 0 | 19 | - | [15 empty] | 1 | [37 empty] | 1 | 1 | only shown for comparison in results
Boudet14 | LP-FILTERING | 0 | 19 | - |  |  |  |  | 1: LP 30Hz (1st order) | [48 empty] | 1 | 1 | only shown for comparison in results
Li13 | L1-CSP | 1 | 15 | - |  |  |  |  | 1: BP 0.5 Hz - 45 Hz | [17 empty] | 2: L1-CSP | [18 empty] | 3: LDA | [11 empty] | 1 | 1
Li13 | L2-CSP | 0 | 15 | - |  |  |  |  | 1: BP 0.5 Hz - 45 Hz | [17 empty] | 2: L2-CSP | [18 empty] | 3: LDA | [11 empty] | 1 | 1
Li13 | TR-CSP | 0 | 15 | - |  |  |  |  | 1: BP 0.5 Hz - 45 Hz | [17 empty] | 2: TR-CSP | [18 empty] | 3: LDA | [11 empty] | 1 | 1
Lawhern13 | DETECT | 1 | 64 | - |  |  |  | 4 * EOG | 1: HP 1Hz (8th order, IIR Butterworth) | [35 empty] | 2: DETECT | [7 empty] | 3 |  |  |  |  | 1 | 1 | The pipeline entered here is not a full description of DETECT, enough details not provided
Nottage13 | Gabor Functions Group 1 | 1 | 64 | - |  |  |  | 2* SMP | 1: LP 200Hz (+Powerline-BP) | [6 empty] | 2: Gabor Functions / Gabor Fitting | [41 empty] | 1 | 0 | Might be interesting to look at gamma band for artifact detection but correction still needs to be done in alpha/beta band for our work
Nottage13 | Gabor Functions Group 2 | 1 | 64 | - |  |  |  | 2* SMP | 1: LP 1000Hz (+Powerline-BP) | [6 empty] | 2: Gabor Functions / Gabor Fitting | [41 empty] | 1 | 0
Yong12 | Ignore | 0 | 15 | [6 empty] | 4 |  | 5: Debounce |  | 1: FFT | 1 | [14 empty] | 2: Stepwise LDA | [15 empty] | 3: LDA | [11 empty] | 0 | 2
Yong12 | Reject | 0 | 15 | [6 empty] | 5 |  | 6: Debounce | 1: 1s sliding with 125ms lag | 2: FFT | 2 | [14 empty] | 3: Stepwise LDA | [15 empty] | 4: LDA | [11 empty] | 0 | 2
Yong12 | SOBI | 0 | 15 | [6 empty] | 6 |  | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | [12 empty] | 2: SOBI |  | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 1 | 2
Yong12 | ERICA | 0 | 15 | [6 empty] | 6 |  | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | [5 empty] | 2: ERICA | [8 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 1 | 2
Yong12 | AMUSE | 0 | 15 | [6 empty] | 6 |  | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | [12 empty] | 2: AMUSE |  | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 1 | 2
Yong12 | SWT-SURE | 0 | 15 | [6 empty] | 6 |  | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | 2: SWT-SURE | [13 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 2 | 2 | global SURE-threshold selection (Steiner's unbiased risk estimate)
Yong12 | ASWTD-Hard | 1 | 15 | [6 empty] | 6 | 2 | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | 2: ASWTD-Hard | [13 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 2 | 2 | adaptive local threshold functions
Yong12 | ASWTD-Soft | 1 | 15 | [6 empty] | 6 | 2 | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | 2: ASWTD-Soft | [13 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 2 | 2
Yong12 | ASWTD-Garrot | 1 | 15 | [6 empty] | 6 | 2 | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | 2: ASWTD-Garrot | [13 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 2 | 2
Yong12 | ASWTD-SBSS | 1 | 15 | [6 empty] | 6 | 2 | 7: Debounce | 1: 1s sliding with 125ms lag | 3: FFT | 3 | 2: ASWTD-SBSS | [13 empty] | 4: Stepwise LDA | [15 empty] | 5: LDA | [11 empty] | 2 | 2
Lau12 | WPLI | 0 | 248 | [5 empty] | 1: HP 1Hz, 4: BP 2-6Hz |  |  |  | 2: 0.5s | 3 | [13 empty] | 5: Weighted Phase Lag Index | [29 empty] | 1 | 2 | Features for channel reduction: std, kurtosis, correlation
Lawhern12 | Autoregression | 1 | 64 | [5 empty] | 1: HP 1Hz |  |  |  | 2: 0.5s | [14 empty] | 3: Auto-regressive model |  |  | 4: SVM (RBF) | [26 empty] | 0 | 2 | only for detection of the movement type
Winkler11 | LPM-Setup1 | 1 | 121 | [5 empty] | 1: HP 2Hz |  |  |  |  | 4: Fingerprint with 13 features | [6 empty] | 3: TDSEP |  | 2 | [33 empty] | 5: LPM (Linear Programming Machine) | 1 | 1
Winkler11 | SVM-Setup1 | 1 | 121 | [5 empty] | 1: HP 2Hz |  |  |  |  | 4: Fingerprint with 13 features | [6 empty] | 3: TDSEP |  | 2 | [29 empty] | 5 |  |  |  |  | 1 | 1
Winkler11 | LDA-Setup1 | 1 | 121 | [5 empty] | 1: HP 2Hz |  |  |  |  | 4: Fingerprint with 13 features | [6 empty] | 3: TDSEP |  | 2 | [22 empty] | 5: LDA | [11 empty] | 1 | 1
Winkler11 | TrainMethod-Setup 2 | 1 | 64 | [5 empty] | 1: BP 0.1-40 Hz |  |  |  |  | 4: Fingerprint with 13 features | [6 empty] | 3: TDSEP |  | 2 | [34 empty] | 1 | 1
Winkler11 | CSP-LDA-Setup 3 | 1 | 90 | [5 empty] | 1: BP 0.1-40 Hz |  |  |  |  | 4: Fingerprint with 13 features | [6 empty] | 3: TDSEP |  | 3 |  |  |  | 5 | [18 empty] | 6: LDA | [11 empty] | 1 | 1
Gwin10 | Original EEG | 0 | 248 | g |  |  |  |  | 1: HP 1Hz | [48 empty] | 0 | 2
Gwin10 | channel-based artifact removal | 0 | 248 | g |  |  |  |  | 1: HP 1Hz |  |  |  | 2: one step | [15 empty] | 3: (from 20 epochs) | [28 empty] | 2 | 0 | for the artifact template, the average of 20 epochs (steps) is used, each epoch is time-warped such that the foot strike happens in the same moment
Gwin10 | IC-based artifact removal | 1 | 248 | g |  |  |  |  | 1: HP 1Hz |  |  |  | 2: one step, 5: one step | [7 empty] | 4: AMICA | [7 empty] | 3: (from 20 epochs), 6: (from 20 epochs) | [28 empty] | 2 | 0
McMenamin10 | ICA | 0 | 128 | [5 empty] | 1: NOTCH 60 Hz, 2: HP 0.5 Hz, 3: BP 8-13Hz | [11 empty] | 5: InfomaxICA |  | 4: PCA (64 ch) | [34 empty] | 0 | 0 | combination of PCA and ICA is not needed anymore
Mosher09 | GSC | 1 |  |  | 102/204 |  |  |  | 1: HP 0.3Hz + LP 333Hz |  |  |  | 4: 7s windows | [9 empty] | 3: Whitening | [10 empty] | 2: SSP (Signal-Space-Projection) |  |  | 5: GSC | [20 empty] | 0 | 2 | Not clear if working for muscle artifacts in EEG, no comparison to other methods
McMenamin09 | LORETA | 1 | 128 | [5 empty] | 1: BP 0.1-250Hz, 2: NOTCH 60Hz | [35 empty] | 3: LORETA | [12 empty] | 0 | 0 | poor method
Delorme04 | - | [58 empty] | no pipeline is needed for this paper, EEGLAB is being explained
Iriarte03 | JADE-ICA | 0 | 64 | - | [16 empty] | 1: JADE | [36 empty] | 1 | 2 | very basic comparison of JADE-ICA VS traditional filters
Iriarte03 | HF_LF_NOTCH | 0 | 64 | - |  |  |  |  | 1: HP (35 Hz; IIR order 4) + LF (0.3 Hz; IIR order 2) + NOTCH (50 Hz) | [48 empty] | 1 | 2
Jung01 | ICA | 0 | 29 | [5 empty] | 1: BP 0.1-50 Hz | [11 empty] | 2 | [36 empty] | 0 | 1 | It is averaged across multiple trials and even persons, therefore not applicable in this form
Summe | 303 | 112 | 299 | 185 | 6 | 0 | 18 | 67 | 259 | 14 | 22 | 18 | 125 | 102 | 29 | 48 | 2 | 2 | 4 | 26 | 109 | 0 | 42 | 2 | 24 | 1 | 7 | 33 | 5 | 13 | 10 | 14 | 2 | 23 | 7 | 2 | 25 | 7 | 1 | 1 | 14 | 18 | 4 | 2 | 13 | 20 | 4 | 7 | 1 | 3 | 3 | 15 | 46 | 5 | 2 | 2 | 11 | 295 | 295

## Results
Unique test setting and pipeline
Reference | Setup ID | Data ID | Metric | Pipeline ID (per paper) | rank in paper | details
Hossain22 | 1 | 1 | SNR | WPD(db1) | 10.5 | rankings were provided per evaluation metric
Hossain22 | 1 | 1 | SNR | WPD(db2) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(db3) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(sym4) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(sym5) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(sym6) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(coif1) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(coif2) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(coif3) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(fk4) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(fk6) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(fk8) | 10.5
Hossain22 | 1 | 1 | SNR | WPD(db1)-CCA | 1
Hossain22 | 1 | 1 | SNR | WPD(db2)-CCA | 2.5
Hossain22 | 1 | 1 | SNR | WPD(db3)-CCA | 10.5
Hossain22 | 1 | 1 | SNR | WPD(fk4)-CCA | 2.5
Hossain22 | 1 | 1 | SNR | WPD(fk6)-CCA | 10.5
Hossain22 | 1 | 1 | SNR | WPD(fk8)-CCA | 18
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db1) | 7
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db2) | 9.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db3) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(sym4) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(sym5) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(sym6) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(coif1) | 9.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(coif2) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(coif3) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk4) | 8
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk6) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk8) | 14.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db1)-CCA | 1
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db2)-CCA | 3
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(db3)-CCA | 4.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk4)-CCA | 2
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk6)-CCA | 4.5
Hossain22 | 2 | 1 | average percentage reduction in artifacts (h) | WPD(fk8)-CCA | 6
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | FastICA-SVM | 3.5 | rankings were provided per evaluation metric, but are the same/similar for every metric
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | FastICA-KNN | 5.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | FastICA-DT | 7.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | FastICA-DA | 1.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | FastICA-NB | 10
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | wICA-SVM | 3.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | wICA-KNN | 5.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | wICA-DT | 7.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | wICA-DA | 1.5
Abu22 | 1 | 1 | Accuracy, Specificity, Sensitivity | wICA-NB | 9
Flo22 | 1 | 1 | SME | APICE | 2
Flo22 | 1 | 1 | SME | APICEa | 4
Flo22 | 1 | 1 | SME | APICE+wICA+iMARA | 2
Flo22 | 1 | 1 | SME | APICE+DSS | 2
Flo22 | 1 | 1 | Ret. epochs | APICE | 2
Flo22 | 1 | 1 | Ret. epochs | APICEa | 4
Flo22 | 1 | 1 | Ret. epochs | APICE+wICA+iMARA | 2
Flo22 | 1 | 1 | Ret. epochs | APICE+DSS | 2
Flo22 | 1 | 2 | SME | APICE | 2.5
Flo22 | 1 | 2 | SME | APICEa | 4
Flo22 | 1 | 2 | SME | APICE+wICA+iMARA | 1
Flo22 | 1 | 2 | SME | APICE+DSS | 2.5
Flo22 | 1 | 2 | Ret. epochs | APICE | 2
Flo22 | 1 | 2 | Ret. epochs | APICEa | 4
Flo22 | 1 | 2 | Ret. epochs | APICE+wICA+iMARA | 2
Flo22 | 1 | 2 | Ret. epochs | APICE+DSS | 2
Flo22 | 2 | 1 | SME | APICE | 1.5
Flo22 | 2 | 1 | SME | Standard Pipeline | 3
Flo22 | 2 | 1 | SME | MADE | 1.5
Flo22 | 2 | 1 | Ret. epochs | APICE | 2
Flo22 | 2 | 1 | Ret. epochs | Standard Pipeline | 1
Flo22 | 2 | 1 | Ret. epochs | MADE | 3
Flo22 | 2 | 2 | SME | APICE | 1
Flo22 | 2 | 2 | SME | Standard Pipeline | 2.5
Flo22 | 2 | 2 | SME | MADE | 2.5
Flo22 | 2 | 2 | Ret. epochs | APICE | 1
Flo22 | 2 | 2 | Ret. epochs | Standard Pipeline | 2.5
Flo22 | 2 | 2 | Ret. epochs | MADE | 2.5
Kumaravel22 | 1 | 1 | PDS | NEAR (ASR R) | 1
Kumaravel22 | 1 | 1 | PDS | NEAR (ASR C) | 2.5
Kumaravel22 | 1 | 1 | PDS | MADE | 4
Kumaravel22 | 1 | 1 | PDS | HAPPE | 2.5
Kumaravel22 | 2 | 2 | accuracy | NEAR (ASR C) | 1
Kumaravel22 | 2 | 2 | accuracy | MADE | 2
Kumaravel22 | 2 | 2 | accuracy | HAPPE | 4
Kumaravel22 | 2 | 2 | accuracy | CRD | 3
Kumaravel22 | 2 | 2 | F1-score | NEAR (ASR C) | 1
Kumaravel22 | 2 | 2 | F1-score | MADE | 2.5
Kumaravel22 | 2 | 2 | F1-score | HAPPE | 4
Kumaravel22 | 2 | 2 | F1-score | CRD | 2.5
Soghoyan21 | 1 | 1 | ROC-AUC | LR | 1.5
Soghoyan21 | 1 | 1 | ROC-AUC | XGB | 3
Soghoyan21 | 1 | 1 | ROC-AUC | SVM | 1.5
Soghoyan21 | 2 | 1 | ROC-AUC | LR | 1
Soghoyan21 | 2 | 1 | ROC-AUC | XGB | 2.5
Soghoyan21 | 2 | 1 | ROC-AUC | SVM | 2.5
Soghoyan21 | 3 | 1 | ROC-AUC | LR | 1.5
Soghoyan21 | 3 | 1 | ROC-AUC | XGB | 1.5
Soghoyan21 | 3 | 1 | ROC-AUC | SVM | 3
Soghoyan21 | 4 | 1 | ROC-AUC | LR | 2
Soghoyan21 | 4 | 1 | ROC-AUC | XGB | 3
Soghoyan21 | 4 | 1 | ROC-AUC | SVM | 1
Soghoyan21 | 5 | 1 | ROC-AUC | LR | 3
Soghoyan21 | 5 | 1 | ROC-AUC | XGB | 2
Soghoyan21 | 5 | 1 | ROC-AUC | SVM | 1
Soghoyan21 | 6 | 1 | ROC-AUC | LR | 1
Soghoyan21 | 6 | 1 | ROC-AUC | XGB | 2
Soghoyan21 | 6 | 1 | ROC-AUC | SVM | 3
Soghoyan21 | 7 | 1 | ROC-AUC | LR | 1
Soghoyan21 | 7 | 1 | ROC-AUC | XGB | 2
Soghoyan21 | 7 | 1 | ROC-AUC | SVM | 3
Soghoyan21 | 1 | 1 | PR-AUC | LR | 2
Soghoyan21 | 1 | 1 | PR-AUC | XGB | 2
Soghoyan21 | 1 | 1 | PR-AUC | SVM | 2
Soghoyan21 | 2 | 1 | PR-AUC | LR | 1.5
Soghoyan21 | 2 | 1 | PR-AUC | XGB | 3
Soghoyan21 | 2 | 1 | PR-AUC | SVM | 1.5
Soghoyan21 | 3 | 1 | PR-AUC | LR | 1
Soghoyan21 | 3 | 1 | PR-AUC | XGB | 3
Soghoyan21 | 3 | 1 | PR-AUC | SVM | 2
Soghoyan21 | 4 | 1 | PR-AUC | LR | 2
Soghoyan21 | 4 | 1 | PR-AUC | XGB | 3
Soghoyan21 | 4 | 1 | PR-AUC | SVM | 1
Soghoyan21 | 5 | 1 | PR-AUC | LR | 2
Soghoyan21 | 5 | 1 | PR-AUC | XGB | 3
Soghoyan21 | 5 | 1 | PR-AUC | SVM | 1
Soghoyan21 | 6 | 1 | PR-AUC | LR | 1
Soghoyan21 | 6 | 1 | PR-AUC | XGB | 2
Soghoyan21 | 6 | 1 | PR-AUC | SVM | 3
Soghoyan21 | 7 | 1 | PR-AUC | LR | 2
Soghoyan21 | 7 | 1 | PR-AUC | XGB | 1
Soghoyan21 | 7 | 1 | PR-AUC | SVM | 3
Soghoyan21 | 1 | 1 | F1-score | LR | 1.5
Soghoyan21 | 1 | 1 | F1-score | XGB | 3
Soghoyan21 | 1 | 1 | F1-score | SVM | 1.5
Soghoyan21 | 2 | 1 | F1-score | LR | 1.5
Soghoyan21 | 2 | 1 | F1-score | XGB | 1.5
Soghoyan21 | 2 | 1 | F1-score | SVM | 3
Soghoyan21 | 3 | 1 | F1-score | LR | 2
Soghoyan21 | 3 | 1 | F1-score | XGB | 1
Soghoyan21 | 3 | 1 | F1-score | SVM | 3
Soghoyan21 | 4 | 1 | F1-score | LR | 3
Soghoyan21 | 4 | 1 | F1-score | XGB | 1
Soghoyan21 | 4 | 1 | F1-score | SVM | 2
Soghoyan21 | 5 | 1 | F1-score | LR | 3
Soghoyan21 | 5 | 1 | F1-score | XGB | 2
Soghoyan21 | 5 | 1 | F1-score | SVM | 1
Soghoyan21 | 6 | 1 | F1-score | LR | 2
Soghoyan21 | 6 | 1 | F1-score | XGB | 1
Soghoyan21 | 6 | 1 | F1-score | SVM | 3
Soghoyan21 | 7 | 1 | F1-score | LR | 2.5
Soghoyan21 | 7 | 1 | F1-score | XGB | 1
Soghoyan21 | 7 | 1 | F1-score | SVM | 2.5
Karpiel21 | 1 | 1 | latency | Original EEG | 3 | no significant difference found
Karpiel21 | 1 | 1 | latency | FIR | 3
Karpiel21 | 1 | 1 | latency | IIR | 3
Karpiel21 | 1 | 1 | latency | FFT | 3
Karpiel21 | 1 | 1 | latency | NOTCH | 3
Karpiel21 | 1 | 1 | amplitude | Original EEG | 1 | the raw signal has the highest amplitude
Karpiel21 | 1 | 1 | amplitude | FIR | 3.5 | no significant difference between filters found
Karpiel21 | 1 | 1 | amplitude | IIR | 3.5
Karpiel21 | 1 | 1 | amplitude | FFT | 3.5
Karpiel21 | 1 | 1 | amplitude | NOTCH | 3.5
Beach21 | 1 | 1 | visual | Original EEG | 2.5 | lower SNR data
Beach21 | 1 | 1 | visual | AF acc | 1
Beach21 | 1 | 1 | visual | AF gyro | 2.5
Beach21 | 1 | 2 | visual | Original EEG | 1 | higher SNR data
Beach21 | 1 | 2 | visual | AF acc | 3
Beach21 | 1 | 2 | visual | AF gyro | 2
Dai21 | 1 | 1 | RRMSE | EEMD-CCA | 1.5
Dai21 | 1 | 1 | RRMSE | sEEMD-CCA | 1.5
Dai21 | 1 | 1 | RRMSE | EEMD-ICA | 3
Dai21 | 1 | 1 | correlation coefficient | EEMD-CCA | 1.5
Dai21 | 1 | 1 | correlation coefficient | sEEMD-CCA | 1.5
Dai21 | 1 | 1 | correlation coefficient | EEMD-ICA | 3
Dai21 | 1 | 1 | runtime | EEMD-CCA | 2
Dai21 | 1 | 1 | runtime | sEEMD-CCA | 1 | runtime is still 2 (vs. 5.5) seconds for a 6s window
Tost21 | 1 | 1 | accuracy | distribution-based double-rejection | 2.5 | no clear comparision between pipelines
Tost21 | 1 | 1 | accuracy | energy double-rejection | 2.5
Tost21 | 1 | 1 | accuracy | hybrid double-rejection | 1
Tost21 | 1 | 1 | F1-score | distribution-based double-rejection | 2.5
Tost21 | 1 | 1 | F1-score | energy double-rejection | 2.5
Tost21 | 1 | 1 | F1-score | hybrid double-rejection | 1
Kaur21 | 1 | 1 | SNR | DWT | 3 | rankings were provided per evaluation metric
Kaur21 | 1 | 1 | SNR | EMD | 3
Kaur21 | 1 | 1 | SNR | EMD-DWT | 3
Kaur21 | 1 | 1 | SNR | EMD-DFA-WPD | 1
Kaur21 | 1 | 1 | MAE | DWT | 3
Kaur21 | 1 | 1 | MAE | EMD | 3
Kaur21 | 1 | 1 | MAE | EMD-DWT | 3
Kaur21 | 1 | 1 | MAE | EMD-DFA-WPD | 1
Rosanne21 | 1 | 1 | accuracy | AF-RF: Original EEG | 6 | SetupID 1 means low physical activity (PA)
Rosanne21 | 1 | 1 | accuracy | AF-RF: ASR | 7
Rosanne21 | 1 | 1 | accuracy | AF-RF: ADJUST | 5
Rosanne21 | 1 | 1 | accuracy | AF-RF: HAPPE | 1
Rosanne21 | 1 | 1 | accuracy | AF-RF: ASR_ADJUST | 9
Rosanne21 | 1 | 1 | accuracy | AF-RF: ASR_wICA | 4
Rosanne21 | 1 | 1 | accuracy | AF-RF: wICA | 2
Rosanne21 | 1 | 1 | accuracy | AF-SVM: Original EEG | 19
Rosanne21 | 1 | 1 | accuracy | AF-SVM: ASR | 23
Rosanne21 | 1 | 1 | accuracy | AF-SVM: ADJUST | 16
Rosanne21 | 1 | 1 | accuracy | AF-SVM: HAPPE | 15
Rosanne21 | 1 | 1 | accuracy | AF-SVM: ASR_ADJUST | 17
Rosanne21 | 1 | 1 | accuracy | AF-SVM: ASR_wICA | 22
Rosanne21 | 1 | 1 | accuracy | AF-SVM: wICA | 20
Rosanne21 | 1 | 1 | accuracy | RF: Original EEG | 14
Rosanne21 | 1 | 1 | accuracy | RF: ASR | 13
Rosanne21 | 1 | 1 | accuracy | RF: ADJUST | 11
Rosanne21 | 1 | 1 | accuracy | RF: HAPPE | 10
Rosanne21 | 1 | 1 | accuracy | RF: ASR_ADJUST | 3
Rosanne21 | 1 | 1 | accuracy | RF: ASR_wICA | 8
Rosanne21 | 1 | 1 | accuracy | RF: wICA | 12
Rosanne21 | 1 | 1 | accuracy | SVM: Original EEG | 28
Rosanne21 | 1 | 1 | accuracy | SVM: ASR | 27
Rosanne21 | 1 | 1 | accuracy | SVM: ADJUST | 24
Rosanne21 | 1 | 1 | accuracy | SVM: HAPPE | 18
Rosanne21 | 1 | 1 | accuracy | SVM: ASR_ADJUST | 21
Rosanne21 | 1 | 1 | accuracy | SVM: ASR_wICA | 25
Rosanne21 | 1 | 1 | accuracy | SVM: wICA | 26
Rosanne21 | 2 | 1 | accuracy | AF-RF: Original EEG | 11 | SetupID 2 means high physical activity (PA)
Rosanne21 | 2 | 1 | accuracy | AF-RF: ASR | 9
Rosanne21 | 2 | 1 | accuracy | AF-RF: ADJUST | 2
Rosanne21 | 2 | 1 | accuracy | AF-RF: HAPPE | 1
Rosanne21 | 2 | 1 | accuracy | AF-RF: ASR_ADJUST | 8
Rosanne21 | 2 | 1 | accuracy | AF-RF: ASR_wICA | 3
Rosanne21 | 2 | 1 | accuracy | AF-RF: wICA | 7
Rosanne21 | 2 | 1 | accuracy | AF-SVM: Original EEG | 17
Rosanne21 | 2 | 1 | accuracy | AF-SVM: ASR | 24
Rosanne21 | 2 | 1 | accuracy | AF-SVM: ADJUST | 18
Rosanne21 | 2 | 1 | accuracy | AF-SVM: HAPPE | 15
Rosanne21 | 2 | 1 | accuracy | AF-SVM: ASR_ADJUST | 22
Rosanne21 | 2 | 1 | accuracy | AF-SVM: ASR_wICA | 23
Rosanne21 | 2 | 1 | accuracy | AF-SVM: wICA | 16
Rosanne21 | 2 | 1 | accuracy | RF: Original EEG | 14
Rosanne21 | 2 | 1 | accuracy | RF: ASR | 12
Rosanne21 | 2 | 1 | accuracy | RF: ADJUST | 6
Rosanne21 | 2 | 1 | accuracy | RF: HAPPE | 5
Rosanne21 | 2 | 1 | accuracy | RF: ASR_ADJUST | 10
Rosanne21 | 2 | 1 | accuracy | RF: ASR_wICA | 4
Rosanne21 | 2 | 1 | accuracy | RF: wICA | 13
Rosanne21 | 2 | 1 | accuracy | SVM: Original EEG | 27
Rosanne21 | 2 | 1 | accuracy | SVM: ASR | 26
Rosanne21 | 2 | 1 | accuracy | SVM: ADJUST | 21
Rosanne21 | 2 | 1 | accuracy | SVM: HAPPE | 19
Rosanne21 | 2 | 1 | accuracy | SVM: ASR_ADJUST | 20
Rosanne21 | 2 | 1 | accuracy | SVM: ASR_wICA | 25
Rosanne21 | 2 | 1 | accuracy | SVM: wICA | 28
Nahmias21 | 1 | 1 | classification accuracy | DCNN | 1 | For motor imagery task
Nahmias21 | 2 | 2 | classification accuracy | DCNN | 1 | For mental arithmetic task
Saba-Sadiya21 | 1 | 1 | f-score | HBOS | 2 | Test setup 1 is when applied to each subject separately, i.e. the models are trained for each subject
Saba-Sadiya21 | 1 | 1 | f-score | LOF | 9
Saba-Sadiya21 | 1 | 1 | f-score | ABOD | 10
Saba-Sadiya21 | 1 | 1 | f-score | OCSVN | 8
Saba-Sadiya21 | 1 | 1 | f-score | LSCP (2*HBOS+OCSVM) | 1
Saba-Sadiya21 | 1 | 1 | f-score | AUTO | 3
Saba-Sadiya21 | 1 | 1 | f-score | PCA | 4
Saba-Sadiya21 | 1 | 1 | f-score | VAE | 5
Saba-Sadiya21 | 1 | 1 | f-score | GAAL | 6
Saba-Sadiya21 | 1 | 1 | f-score | MGAAL | 7
Saba-Sadiya21 | 1 | 1 | Cohen's K | HBOS | 2
Saba-Sadiya21 | 1 | 1 | Cohen's K | LOF | 9
Saba-Sadiya21 | 1 | 1 | Cohen's K | ABOD | 10
Saba-Sadiya21 | 1 | 1 | Cohen's K | OCSVN | 8
Saba-Sadiya21 | 1 | 1 | Cohen's K | LSCP (2*HBOS+OCSVM) | 1
Saba-Sadiya21 | 1 | 1 | Cohen's K | AUTO | 3
Saba-Sadiya21 | 1 | 1 | Cohen's K | PCA | 4
Saba-Sadiya21 | 1 | 1 | Cohen's K | VAE | 5
Saba-Sadiya21 | 1 | 1 | Cohen's K | GAAL | 6
Saba-Sadiya21 | 1 | 1 | Cohen's K | MGAAL | 7
Saba-Sadiya21 | 1 | 2 | f-score | HBOS | 3
Saba-Sadiya21 | 1 | 2 | f-score | LOF | 8.5
Saba-Sadiya21 | 1 | 2 | f-score | ABOD | 10
Saba-Sadiya21 | 1 | 2 | f-score | OCSVN | 7
Saba-Sadiya21 | 1 | 2 | f-score | LSCP (2*HBOS+OCSVM) | 1.5
Saba-Sadiya21 | 1 | 2 | f-score | AUTO | 1.5
Saba-Sadiya21 | 1 | 2 | f-score | PCA | 5
Saba-Sadiya21 | 1 | 2 | f-score | VAE | 4
Saba-Sadiya21 | 1 | 2 | f-score | GAAL | 8.5
Saba-Sadiya21 | 1 | 2 | f-score | MGAAL | 6
Saba-Sadiya21 | 1 | 2 | Cohen's K | HBOS | 3
Saba-Sadiya21 | 1 | 2 | Cohen's K | LOF | 8
Saba-Sadiya21 | 1 | 2 | Cohen's K | ABOD | 10
Saba-Sadiya21 | 1 | 2 | Cohen's K | OCSVN | 7
Saba-Sadiya21 | 1 | 2 | Cohen's K | LSCP (2*HBOS+OCSVM) | 2
Saba-Sadiya21 | 1 | 2 | Cohen's K | AUTO | 1
Saba-Sadiya21 | 1 | 2 | Cohen's K | PCA | 5
Saba-Sadiya21 | 1 | 2 | Cohen's K | VAE | 4
Saba-Sadiya21 | 1 | 2 | Cohen's K | GAAL | 9
Saba-Sadiya21 | 1 | 2 | Cohen's K | MGAAL | 6
Saba-Sadiya21 | 2 | 1 | f-score | HBOS | 2 | Test setup 1 is when applied to all subject, i.e. the models are trained only once per dataset
Saba-Sadiya21 | 2 | 1 | f-score | LOF | 9
Saba-Sadiya21 | 2 | 1 | f-score | ABOD | 10
Saba-Sadiya21 | 2 | 1 | f-score | OCSVN | 8
Saba-Sadiya21 | 2 | 1 | f-score | LSCP (2*HBOS+OCSVM) | 1
Saba-Sadiya21 | 2 | 1 | f-score | AUTO | 3
Saba-Sadiya21 | 2 | 1 | f-score | PCA | 4
Saba-Sadiya21 | 2 | 1 | f-score | VAE | 5
Saba-Sadiya21 | 2 | 1 | f-score | GAAL | 7
Saba-Sadiya21 | 2 | 1 | f-score | MGAAL | 6
Saba-Sadiya21 | 2 | 1 | Cohen's K | HBOS | 2
Saba-Sadiya21 | 2 | 1 | Cohen's K | LOF | 9
Saba-Sadiya21 | 2 | 1 | Cohen's K | ABOD | 10
Saba-Sadiya21 | 2 | 1 | Cohen's K | OCSVN | 8
Saba-Sadiya21 | 2 | 1 | Cohen's K | LSCP (2*HBOS+OCSVM) | 1
Saba-Sadiya21 | 2 | 1 | Cohen's K | AUTO | 3.5
Saba-Sadiya21 | 2 | 1 | Cohen's K | PCA | 3.5
Saba-Sadiya21 | 2 | 1 | Cohen's K | VAE | 5
Saba-Sadiya21 | 2 | 1 | Cohen's K | GAAL | 7
Saba-Sadiya21 | 2 | 1 | Cohen's K | MGAAL | 6
Saba-Sadiya21 | 2 | 2 | f-score | HBOS | 1
Saba-Sadiya21 | 2 | 2 | f-score | LOF | 8
Saba-Sadiya21 | 2 | 2 | f-score | ABOD | 10
Saba-Sadiya21 | 2 | 2 | f-score | OCSVN | 7
Saba-Sadiya21 | 2 | 2 | f-score | LSCP (2*HBOS+OCSVM) | 2
Saba-Sadiya21 | 2 | 2 | f-score | AUTO | 5
Saba-Sadiya21 | 2 | 2 | f-score | PCA | 3
Saba-Sadiya21 | 2 | 2 | f-score | VAE | 4
Saba-Sadiya21 | 2 | 2 | f-score | GAAL | 9
Saba-Sadiya21 | 2 | 2 | f-score | MGAAL | 6
Saba-Sadiya21 | 2 | 2 | Cohen's K | HBOS | 1
Saba-Sadiya21 | 2 | 2 | Cohen's K | LOF | 8
Saba-Sadiya21 | 2 | 2 | Cohen's K | ABOD | 10
Saba-Sadiya21 | 2 | 2 | Cohen's K | OCSVN | 7
Saba-Sadiya21 | 2 | 2 | Cohen's K | LSCP (2*HBOS+OCSVM) | 2
Saba-Sadiya21 | 2 | 2 | Cohen's K | AUTO | 5
Saba-Sadiya21 | 2 | 2 | Cohen's K | PCA | 3
Saba-Sadiya21 | 2 | 2 | Cohen's K | VAE | 4
Saba-Sadiya21 | 2 | 2 | Cohen's K | GAAL | 9
Saba-Sadiya21 | 2 | 2 | Cohen's K | MGAAL | 6
Saba-Sadiya21 | 3 | 3 | Accuracy | LSCP (2*HBOS+OCSVM) | 1 | Here, a SVM classifies for the artifact-corrected and original EEG
Saba-Sadiya21 | 3 | 3 | Accuracy | Original EEG | 2
Saba-Sadiya21 | 4 | 3 | Accuracy | LSCP (2*HBOS+OCSVM) | 1 | Two trials, setups with All trails and rejected trials only
Saba-Sadiya21 | 4 | 3 | Accuracy | Original EEG | 2
Li21 | 1 | 1 | z-scored power | Original EEG | 4 | We only added the real data results here, ignoring simulated results
Li21 | 1 | 1 | z-scored power | ERASE-Real-EMG | 1.5
Li21 | 1 | 1 | z-scored power | ERASE-Simulated-EMG | 1.5
Li21 | 1 | 1 | z-scored power | Conventional-ICA | 3
Anders20 | - | - | - | - | -
Zhang20 | - | - | - | - | - | results do not show wether the presented electrode system is capable of removing real artifacts in real settings better than estiablished systems - further work suggested for testing
Leach20 | 1 | 1 | Percent Agreement | adjusted ADJUST | 1.5 | Data ID 1: adult, 2: children, 3: infants
Leach20 | 1 | 1 | Percent Agreement | original ADJUST | 3
Leach20 | 1 | 1 | Percent Agreement | ICLabel | 1.5
Leach20 | 1 | 2 | Percent Agreement | adjusted ADJUST | 1
Leach20 | 1 | 2 | Percent Agreement | original ADJUST | 3
Leach20 | 1 | 2 | Percent Agreement | ICLabel | 2
Leach20 | 1 | 3 | Percent Agreement | adjusted ADJUST | 1.5
Leach20 | 1 | 3 | Percent Agreement | original ADJUST | 1.5
Leach20 | 1 | 3 | Percent Agreement | ICLabel | 3
Leach20 | 1 | 1 | Trails retained | adjusted ADJUST | 1.5
Leach20 | 1 | 1 | Trails retained | original ADJUST | 3
Leach20 | 1 | 1 | Trails retained | ICLabel | 1.5
Leach20 | 1 | 1 | Trails retained | No correction | 4
Leach20 | 1 | 2 | Trails retained | adjusted ADJUST | 1
Leach20 | 1 | 2 | Trails retained | original ADJUST | 3
Leach20 | 1 | 2 | Trails retained | ICLabel | 2
Leach20 | 1 | 2 | Trails retained | No correction | 4
Leach20 | 1 | 3 | Trails retained | adjusted ADJUST | 1.5
Leach20 | 1 | 3 | Trails retained | original ADJUST | 3.5
Leach20 | 1 | 3 | Trails retained | ICLabel | 1.5
Leach20 | 1 | 3 | Trails retained | No correction | 3.5
Leach20 | 1 | 1 | Reliability | adjusted ADJUST | 2
Leach20 | 1 | 1 | Reliability | original ADJUST | 2
Leach20 | 1 | 1 | Reliability | ICLabel | 2
Leach20 | 1 | 1 | Reliability | No correction | 4
Leach20 | 1 | 2 | Reliability | adjusted ADJUST | 1.5
Leach20 | 1 | 2 | Reliability | original ADJUST | 3.5
Leach20 | 1 | 2 | Reliability | ICLabel | 1.5
Leach20 | 1 | 2 | Reliability | No correction | 3.5
Leach20 | 1 | 3 | Reliability | adjusted ADJUST | 1
Leach20 | 1 | 3 | Reliability | original ADJUST | 2.5
Leach20 | 1 | 3 | Reliability | ICLabel | 2.5
Leach20 | 1 | 3 | Reliability | No correction | 4
Saini20 | 1 | 1 | PRD | EEMD-CCA | 4
Saini20 | 1 | 1 | PRD | EEMD | 2
Saini20 | 1 | 1 | PRD | EMD-CCA | 6
Saini20 | 1 | 1 | PRD | Wavelet Denoising | 5
Saini20 | 1 | 1 | PRD | EEMD-MCCA | 3
Saini20 | 1 | 1 | PRD | VMD-cross | 1
Saini20 | 2 | 2 | PRD | EEMD-CCA | 4
Saini20 | 2 | 2 | PRD | EEMD | 2
Saini20 | 2 | 2 | PRD | EMD-CCA | 6
Saini20 | 2 | 2 | PRD | Wavelet Denoising | 5
Saini20 | 2 | 2 | PRD | EEMD-MCCA | 3
Saini20 | 2 | 2 | PRD | VMD-cross | 1
Saini20 | 3 | 3 | PRD | EEMD-CCA | 4
Saini20 | 3 | 3 | PRD | EEMD | 2
Saini20 | 3 | 3 | PRD | EMD-CCA | 6
Saini20 | 3 | 3 | PRD | Wavelet Denoising | 5
Saini20 | 3 | 3 | PRD | EEMD-MCCA | 3
Saini20 | 3 | 3 | PRD | VMD-cross | 1
Saini20 | 1 | 1 | RMSE | EEMD-CCA | 3.5
Saini20 | 1 | 1 | RMSE | EEMD | 3.5
Saini20 | 1 | 1 | RMSE | EMD-CCA | 6
Saini20 | 1 | 1 | RMSE | Wavelet Denoising | 3.5
Saini20 | 1 | 1 | RMSE | EEMD-MCCA | 3.5
Saini20 | 1 | 1 | RMSE | VMD-cross | 1
Saini20 | 2 | 2 | RMSE | EEMD-CCA | 4
Saini20 | 2 | 2 | RMSE | EEMD | 2
Saini20 | 2 | 2 | RMSE | EMD-CCA | 6
Saini20 | 2 | 2 | RMSE | Wavelet Denoising | 4
Saini20 | 2 | 2 | RMSE | EEMD-MCCA | 4
Saini20 | 2 | 2 | RMSE | VMD-cross | 1
Saini20 | 3 | 3 | RMSE | EEMD-CCA | 3.5
Saini20 | 3 | 3 | RMSE | EEMD | 3.5
Saini20 | 3 | 3 | RMSE | EMD-CCA | 6
Saini20 | 3 | 3 | RMSE | Wavelet Denoising | 3.5
Saini20 | 3 | 3 | RMSE | EEMD-MCCA | 3.5
Saini20 | 3 | 3 | RMSE | VMD-cross | 1
Mucarquer20 | 1 | 1 | RRMSE | SuBAR | 4.5 | only evaluated real EEG data with real EMG data
Mucarquer20 | 1 | 1 | RRMSE | EEMD-CCA | 4.5
Mucarquer20 | 1 | 1 | RRMSE | EEMD-CCA-2EMG | 2.5
Mucarquer20 | 1 | 1 | RRMSE | EEMD-CCA-4EMG | 2.5
Mucarquer20 | 1 | 1 | RRMSE | EEMD-CCA-8EMG | 1
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | Single-Layer-AMICA | 5 | results evaluated: Relative power (db), scalp map, dipole locations and ERSP plots
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | Dual-Layer-AMICA | 5
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | Noise-Cancel-AMICA | 5
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | ASR-AMICA | 2
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | ASR-merged-AMICA | 2
Nordin20 | 1 | 1 | Relative power (db), scalp map, dipole locations and ERSP plots | PCA-CCA-merged-AMICA | 2
Liu19 | 1 | 1 | RRMSE | CCA | 3
Liu19 | 1 | 1 | RRMSE | EEMD-CCA | 2
Liu19 | 1 | 1 | RRMSE | SSH-CCA | 1
Liu19 | 1 | 1 | accuracy | CCA | 3
Liu19 | 1 | 1 | accuracy | EEMD-CCA | 2
Liu19 | 1 | 1 | accuracy | SSH-CCA | 1
Liu19 | 1 | 2 | PDS | CCA | 3
Liu19 | 1 | 2 | PDS | EEMD-CCA | 1.5
Liu19 | 1 | 2 | PDS | SSH-CCA | 1.5
Liu19 | 1 | 2 | visual | CCA | 3
Liu19 | 1 | 2 | visual | EEMD-CCA | 2
Liu19 | 1 | 2 | visual | SSH-CCA | 1
Benda19 | 1 | 1 | peak detection efficacy | MC-Wiener-Filtered | 1 | results here can only be compared by metrics and applied multi-channel Wiener filtering
Benda19 | 1 | 1 | peak detection efficacy | MC-Wiener-Unfiltered | 2
Benda19 | 2 | 1 | peak detection accuracy | MC-Wiener-Filtered | 1.5
Benda19 | 2 | 1 | peak detection accuracy | MC-Wiener-Unfiltered | 1.5
Val-Calvo19 | 1 | 1 | RMSE | EAWICA-ALL | 3
Val-Calvo19 | 1 | 1 | RMSE | ICA-W-ALL | 2
Val-Calvo19 | 1 | 1 | RMSE | EAWICA-DELTA | 4
Val-Calvo19 | 1 | 1 | RMSE | ICA-W-DELTA | 1
Val-Calvo19 | 1 | 1 | Correlation | EAWICA-ALL | 3
Val-Calvo19 | 1 | 1 | Correlation | ICA-W-ALL | 2
Val-Calvo19 | 1 | 1 | Correlation | EAWICA-DELTA | 4
Val-Calvo19 | 1 | 1 | Correlation | ICA-W-DELTA | 1
Val-Calvo19 | 1 | 1 | Mutual Information | EAWICA-ALL | 3
Val-Calvo19 | 1 | 1 | Mutual Information | ICA-W-ALL | 2
Val-Calvo19 | 1 | 1 | Mutual Information | EAWICA-DELTA | 4
Val-Calvo19 | 1 | 1 | Mutual Information | ICA-W-DELTA | 1
Val-Calvo19 | 1 | 1 | Time | EAWICA-ALL | 1
Val-Calvo19 | 1 | 1 | Time | ICA-W-ALL | 3
Val-Calvo19 | 1 | 1 | Time | EAWICA-DELTA | 2
Val-Calvo19 | 1 | 1 | Time | ICA-W-DELTA | 4
Val-Calvo19 | 2 | 1 | Accuracy | Linear dynamic system | 2
Val-Calvo19 | 2 | 1 | Accuracy | Savetzky-Golay | 1
Val-Calvo19 | 2 | 1 | Time | Linear dynamic system | 2
Val-Calvo19 | 2 | 1 | Time | Savetzky-Golay | 1
Saavedra19 | 1 | 1 | Accuracy | xDAWN | 3 | Setup-ID here is the number of repititions (evolution of algorithm upon repitition)
Saavedra19 | 1 | 1 | Accuracy | METS_SEWS_1 | 1
Saavedra19 | 1 | 1 | Accuracy | METS_SEWS_2 | 2
Saavedra19 | 2 | 1 | Accuracy | xDAWN | 2
Saavedra19 | 2 | 1 | Accuracy | METS_SEWS_1 | 3
Saavedra19 | 2 | 1 | Accuracy | METS_SEWS_2 | 1
Saavedra19 | 3 | 1 | Accuracy | xDAWN | 3
Saavedra19 | 3 | 1 | Accuracy | METS_SEWS_1 | 2
Saavedra19 | 3 | 1 | Accuracy | METS_SEWS_2 | 1
Saavedra19 | 4 | 1 | Accuracy | xDAWN | 1
Saavedra19 | 4 | 1 | Accuracy | METS_SEWS_1 | 3
Saavedra19 | 4 | 1 | Accuracy | METS_SEWS_2 | 2
Saavedra19 | 5 | 1 | Accuracy | xDAWN | 1
Saavedra19 | 5 | 1 | Accuracy | METS_SEWS_1 | 3
Saavedra19 | 5 | 1 | Accuracy | METS_SEWS_2 | 2
Mur19 | 1 | 1 | visual | UAR | 1.5 | UAR has too strong assumption to be applicable
Mur19 | 1 | 1 | visual | MARA | 3
Mur19 | 1 | 1 | visual | SASICA | 1.5
Mur19 | 1 | 1 | visual | wICA | 4
Blum19 | 1 | 1 | Specificity | Original EEG | 2
Blum19 | 1 | 1 | Specificity | ASR | 2
Blum19 | 1 | 1 | Specificity | rASR | 2
Blum19 | 1 | 1 | Sensitivity | ASR | 1.5
Blum19 | 1 | 1 | Sensitivity | rASR | 1.5
Blum19 | 1 | 1 | Latency | ASR | 2
Blum19 | 1 | 1 | Latency | rASR | 1 | rASR is about 45% faster
Nordin19 | 1 | 1 | spectral perturbation | Infomax | 2 | Single Layer System Measurement
Nordin19 | 1 | 1 | spectral perturbation | Infomax | 1 | Dual Layer System Measurement
Grosselin19 | 1 | 1 | Accuracy | LDA | 4 | Only artBA (id=1) and artMM (id=2) databases evaluated for classifier performance
Grosselin19 | 1 | 1 | Accuracy | Linear SVM | 2
Grosselin19 | 1 | 1 | Accuracy | Euclidean | 3
Grosselin19 | 1 | 1 | Accuracy | Weighted kNNs | 1
Grosselin19 | 1 | 1 | AUC | LDA | 2.5
Grosselin19 | 1 | 1 | AUC | Linear SVM | 2.5
Grosselin19 | 1 | 1 | AUC | Euclidean | 2.5
Grosselin19 | 1 | 1 | AUC | Weighted kNNs | 2.5
Grosselin19 | 1 | 2 | Accuracy | LDA | 4
Grosselin19 | 1 | 2 | Accuracy | Linear SVM | 1
Grosselin19 | 1 | 2 | Accuracy | Euclidean | 2.5
Grosselin19 | 1 | 2 | Accuracy | Weighted kNNs | 2.5
Grosselin19 | 1 | 2 | AUC | LDA | 4
Grosselin19 | 1 | 2 | AUC | Linear SVM | 2
Grosselin19 | 1 | 2 | AUC | Euclidean | 2
Grosselin19 | 1 | 2 | AUC | Weighted kNNs | 2
Li18 | 1 | 1 | accuracy | MEMD+SampEn | 1 | not compared with others
Chu18 | 1 | 1 | accuracy | LSP-DBN | 1 | setupID here is used to distinguish between DBN and SVM pipeline results, as comparison results of SVM and DBN could not be listed in detail
Chu18 | 1 | 1 | accuracy | FFT-DBN | 2.5
Chu18 | 1 | 1 | accuracy | Welch-DBN | 2.5
Chu18 | 2 | 1 | accuracy | LSP-SVM | 1
Chu18 | 2 | 1 | accuracy | FFT-SVM | 2.5
Chu18 | 2 | 1 | accuracy | Welch-SVM | 2.5
San-Martin18 | - | - | - | -
Sebek18 | - | - | - | - | - | Other pipelines not explained in detail, therefore skipped comparison here. Work would benefit from proposed method combined with different ICA approaches
Arad18 | 1 | 1 | group average, SD | AMICA | 1.5 | Compared to already previously existing ASR strategy at various speeds
Arad18 | 1 | 1 | group average, SD | ASR | 1.5
Symeonidou18 | - | - | - | - | - | Info: No comparison of results is needed for this paper
Stone18 | - | - | - | - | - | Info: Method was not compared against others
Frolich18 | 1 | 1 | ERD peak score | Original EEG | 9
Frolich18 | 1 | 1 | ERD peak score | TDSEP | 7
Frolich18 | 1 | 1 | ERD peak score | Infomax | 8
Frolich18 | 1 | 1 | ERD peak score | FastICA | 6
Frolich18 | 1 | 1 | ERD peak score | HP-TDSEP | 5
Frolich18 | 1 | 1 | ERD peak score | HP-Infomax | 1
Frolich18 | 1 | 1 | ERD peak score | HP-FastICA | 4
Frolich18 | 1 | 1 | ERD peak score | SSD | 3
Frolich18 | 1 | 1 | ERD peak score | Fourier-ICA | 2
Frolich18 | 1 | 1 | dipolarity | Original EEG | 9
Frolich18 | 1 | 1 | dipolarity | TDSEP | 3
Frolich18 | 1 | 1 | dipolarity | Infomax | 1
Frolich18 | 1 | 1 | dipolarity | FastICA | 4
Frolich18 | 1 | 1 | dipolarity | HP-TDSEP | 7
Frolich18 | 1 | 1 | dipolarity | HP-Infomax | 2
Frolich18 | 1 | 1 | dipolarity | HP-FastICA | 6
Frolich18 | 1 | 1 | dipolarity | SSD | 8
Frolich18 | 1 | 1 | dipolarity | Fourier-ICA | 5
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | HAPPE | 1 | HAPPE introduced here and compared to others methods; only entered results from Fig.9 here
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | ICA-MARA | 7
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | Manual | 3.5
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | ASR | 3.5
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | ADJUST | 3.5
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | SASICA | 3.5
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | FASTER | 7
Gabard-Durnam18 | 1 | 1 | average percent of artifact level retained, average percent of EEG signal retained | FASTER-MARA | 7
Tamburro18 | 1 | 1 | Accuracy | Gel, 20 IC | 6 | data ID 1 belongs to myogenic artifacty only
Tamburro18 | 1 | 1 | Accuracy | Gel, 50 IC | 3
Tamburro18 | 1 | 1 | Accuracy | Gel, 80 IC | 3
Tamburro18 | 1 | 1 | Accuracy | Dry, 20 IC | 5
Tamburro18 | 1 | 1 | Accuracy | Dry, 50 IC | 3
Tamburro18 | 1 | 1 | Accuracy | Dry, 80 IC | 1
Tamburro18 | 1 | 1 | FOR | Gel, 20 IC | 6
Tamburro18 | 1 | 1 | FOR | Gel, 50 IC | 4.5
Tamburro18 | 1 | 1 | FOR | Gel, 80 IC | 1
Tamburro18 | 1 | 1 | FOR | Dry, 20 IC | 4.5
Tamburro18 | 1 | 1 | FOR | Dry, 50 IC | 3
Tamburro18 | 1 | 1 | FOR | Dry, 80 IC | 2
Tamburro18 | 1 | 1 | sensitivity | Gel, 20 IC | 2
Tamburro18 | 1 | 1 | sensitivity | Gel, 50 IC | 2
Tamburro18 | 1 | 1 | sensitivity | Gel, 80 IC | 2
Tamburro18 | 1 | 1 | sensitivity | Dry, 20 IC | 4
Tamburro18 | 1 | 1 | sensitivity | Dry, 50 IC | 6
Tamburro18 | 1 | 1 | sensitivity | Dry, 80 IC | 5
Tamburro18 | 2 | 1 | SNR | Gel, 20 IC | 6
Tamburro18 | 2 | 1 | SNR | Gel, 50 IC | 4
Tamburro18 | 2 | 1 | SNR | Gel, 80 IC | 4
Tamburro18 | 2 | 1 | SNR | Gel, Original EEG | 8
Tamburro18 | 2 | 1 | SNR | Dry, 20 IC | 4
Tamburro18 | 2 | 1 | SNR | Dry, 50 IC | 1.5
Tamburro18 | 2 | 1 | SNR | Dry, 80 IC | 1.5
Tamburro18 | 2 | 1 | SNR | Dry, Original EEG | 7 | unprocessed dry electrodes are here less affected by artifacts, therefore lack of comparability
Lin18 | 1 | 1 | visual | CCA+GMM | 1.5 | no real evaluation metric is used
Lin18 | 1 | 1 | visual | ASR | 1.5 | results are not really compared
Delisle-Rodriguez17 | 1 | 1 | Attenuation(dB) | Filter Ad WAR | 1 | check the order
Delisle-Rodriguez17 | 1 | 1 | Attenuation(dB) | Filter WAR | 2
Delisle-Rodriguez17 | 1 | 1 | SNR(dB) | Filter Ad WAR | 2
Delisle-Rodriguez17 | 1 | 1 | SNR(dB) | Filter WAR | 1
Delisle-Rodriguez17 | 1 | 1 | Coherence(dB) | Filter Ad WAR | 1.5
Delisle-Rodriguez17 | 1 | 1 | Coherence(dB) | Filter WAR | 1.5
Delisle-Rodriguez17 | 1 | 1 | FPR(%) | Filter Ad WAR | 2
Delisle-Rodriguez17 | 1 | 1 | FPR(%) | Filter WAR | 1
Delisle-Rodriguez17 | 1 | 1 | accuracy | Filter Ad WAR | 2 | The WAR filter is here really better
Delisle-Rodriguez17 | 1 | 1 | accuracy | Filter WAR | 1
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | Ad LAR | 2.5 | Signifiance levels are provided for AdLAR and AdWAR to each other respectively. Therefore for each of them there are two setups listed
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | Ad WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | Riemann Kernel | 5
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | CSP | 6.5
Delisle-Rodriguez17 | 2 | 2 | TPR(%) | Filter-Bank CSP | 6.5
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | Ad WAR | 2
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | LAR | 2
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | WAR | 2
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | Riemann Kernel | 4
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | CSP | 5.5
Delisle-Rodriguez17 | 3 | 2 | TPR(%) | Filter-Bank CSP | 5.5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | Ad LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | Ad WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | Riemann Kernel | 5
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | CSP | 7
Delisle-Rodriguez17 | 2 | 2 | FPR(%) | Filter-Bank CSP | 6
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | Ad WAR | 1
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | LAR | 3
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | WAR | 3
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | Riemann Kernel | 3
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | CSP | 6
Delisle-Rodriguez17 | 3 | 2 | FPR(%) | Filter-Bank CSP | 5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | Ad LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | Ad WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | Riemann Kernel | 5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | CSP | 6.5
Delisle-Rodriguez17 | 2 | 2 | F1(%) | Filter-Bank CSP | 6.5
Delisle-Rodriguez17 | 3 | 2 | F1(%) | Ad WAR | 1
Delisle-Rodriguez17 | 3 | 2 | F1(%) | LAR | 2.5
Delisle-Rodriguez17 | 3 | 2 | F1(%) | WAR | 2.5
Delisle-Rodriguez17 | 3 | 2 | F1(%) | Riemann Kernel | 4
Delisle-Rodriguez17 | 3 | 2 | F1(%) | CSP | 5.5
Delisle-Rodriguez17 | 3 | 2 | F1(%) | Filter-Bank CSP | 5.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | Ad LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | Ad WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | LAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | WAR | 2.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | Riemann Kernel | 5
Delisle-Rodriguez17 | 2 | 2 | accuracy | CSP | 6.5
Delisle-Rodriguez17 | 2 | 2 | accuracy | Filter-Bank CSP | 6.5
Delisle-Rodriguez17 | 3 | 2 | accuracy | Ad WAR | 1
Delisle-Rodriguez17 | 3 | 2 | accuracy | LAR | 2.5
Delisle-Rodriguez17 | 3 | 2 | accuracy | WAR | 2.5
Delisle-Rodriguez17 | 3 | 2 | accuracy | Riemann Kernel | 4
Delisle-Rodriguez17 | 3 | 2 | accuracy | CSP | 5.5
Delisle-Rodriguez17 | 3 | 2 | accuracy | Filter-Bank CSP | 5.5
Wang17 | - | - | - | - | -
Jas17 | 1 | 1 | infinity norm | Autoreject | 2.5 | comparison was done in so many diverse ways that no significant rank could be applied
Jas17 | 1 | 1 | infinity norm | FASTER | 2.5
Jas17 | 1 | 1 | infinity norm | SNS | 2.5
Jas17 | 1 | 1 | infinity norm | RANSAC | 2.5
Jas17 | 1 | 1 | infinity norm | NoReject | 5
Oliveira17 | 1 | 1 | spectral power | CR1 | 1.5 | There is no significant difference between the traditional (CR1) and template correlation rejection (TCR) method
Oliveira17 | 1 | 1 | spectral power | TCR | 1.5
Oliveira17 | 2 | 1 | spectral power | CR1 | 1.5
Oliveira17 | 2 | 1 | spectral power | TCR | 1.5
Oliveira17 | 3 | 1 | spectral power | CR1 | 1.5
Oliveira17 | 3 | 1 | spectral power | TCR | 1.5
Oliveira17 | 4 | 1 | spectral power | CR1 | 1.5
Oliveira17 | 4 | 1 | spectral power | TCR | 1.5
Oliveira17 | 5 | 1 | spectral power | CR1 | 1.5
Oliveira17 | 5 | 1 | spectral power | TCR | 1.5
DelPozo-Banos17 | 1 | 1 | TAR | None-WithLCF | 5 | comparison per Metric given; only compared results from real data
DelPozo-Banos17 | 1 | 1 | TAR | ADJUST-WithLCF | 1
DelPozo-Banos17 | 1 | 1 | TAR | FASTER-WithLCF | 2.5
DelPozo-Banos17 | 1 | 1 | TAR | ICAW-WithLCF | 4
DelPozo-Banos17 | 1 | 1 | TAR | None-WithoutLCF | 8
DelPozo-Banos17 | 1 | 1 | TAR | ADJUST-WithoutLCF | 2.5
DelPozo-Banos17 | 1 | 1 | TAR | FASTER-WithoutLCF | 6
DelPozo-Banos17 | 1 | 1 | TAR | ICAW-WithoutLCF | 7
DelPozo-Banos17 | 1 | 1 | FAR | None-WithLCF | 7
DelPozo-Banos17 | 1 | 1 | FAR | ADJUST-WithLCF | 2
DelPozo-Banos17 | 1 | 1 | FAR | FASTER-WithLCF | 5.5
DelPozo-Banos17 | 1 | 1 | FAR | ICAW-WithLCF | 2
DelPozo-Banos17 | 1 | 1 | FAR | None-WithoutLCF | 8
DelPozo-Banos17 | 1 | 1 | FAR | ADJUST-WithoutLCF | 4
DelPozo-Banos17 | 1 | 1 | FAR | FASTER-WithoutLCF | 5.5
DelPozo-Banos17 | 1 | 1 | FAR | ICAW-WithoutLCF | 2
Weiss17 | - | - | - | - | - | AR1 is not described with enough detail in the paper to make a meaningful comparison here
Roy17 | 1 | 1 | RMSE | EEMD-CCA-DWT | 4
Roy17 | 1 | 1 | RMSE | EEMD-GECCA-DWT | 2
Roy17 | 1 | 1 | RMSE | EEMD-CCA-SWT | 3
Roy17 | 1 | 1 | RMSE | EEMD-GECCA-SWT | 1
Roy17 | 1 | 1 | Correlation | EEMD-CCA-DWT | 4
Roy17 | 1 | 1 | Correlation | EEMD-GECCA-DWT | 1
Roy17 | 1 | 1 | Correlation | EEMD-CCA-SWT | 3
Roy17 | 1 | 1 | Correlation | EEMD-GECCA-SWT | 2
Roy17 | 1 | 1 | Lambda | EEMD-CCA-DWT | 3
Roy17 | 1 | 1 | Lambda | EEMD-GECCA-DWT | 2
Roy17 | 1 | 1 | Lambda | EEMD-CCA-SWT | 4
Roy17 | 1 | 1 | Lambda | EEMD-GECCA-SWT | 1
Roy17 | 1 | 1 | SNR | EEMD-CCA-DWT | 3
Roy17 | 1 | 1 | SNR | EEMD-GECCA-DWT | 2
Roy17 | 1 | 1 | SNR | EEMD-CCA-SWT | 4
Roy17 | 1 | 1 | SNR | EEMD-GECCA-SWT | 1
Melman16 | 1 | 1 | Power density spectrum | Robust Method w/Bayes | 1.5
Melman16 | 1 | 1 | Power density spectrum | Standard Method w/Jackknife | 1.5
Melman16 | 1 | 2 | Power density spectrum | Robust Method w/Bayes | 1
Melman16 | 1 | 2 | Power density spectrum | Standard Method w/Jackknife | 2
Melman16 | 1 | 3 | Power density spectrum | Robust Method w/Bayes | 1
Melman16 | 1 | 3 | Power density spectrum | Standard Method w/Jackknife | 2
Oliveira16 | - | - | - | -
Cheveigne16 | 1 | 1 | visual | STAR | 1
Cheveigne16 | 1 | 1 | visual | Original EEG | 2 | Star performs better than nothing
Cheveigne16 | 2 | 1 | visual | STAR | 2.5
Cheveigne16 | 2 | 1 | visual | JD | 2.5
Cheveigne16 | 2 | 1 | visual | STAR+JD | 1
Cheveigne16 | 2 | 1 | visual | Original EEG | 4
Costa16 | - | - | - | -
Nathan16 | - | - | - | -
Zou16 | 1 | 1 | accuracy | Our Method | 1 | they all use an ICA with 8 channels only, therefore might not be representative
Zou16 | 1 | 1 | accuracy | FASTER | 2.5
Zou16 | 1 | 1 | accuracy | ADJUST | 2.5
Zou16 | 1 | 1 | accuracy | Auto-mutual information | 4.5 | these methods are not built to detect non-biological artifacts, therefore a bad comparison
Zou16 | 1 | 1 | accuracy | K-means with similarity | 4.5
Gordon15 | 1 | 1 | visual | IMICA | 2.5 | entered for both datasets median, only muscle artifact results are compared here (Fig. 4)
Gordon15 | 1 | 1 | visual | Infomax | 2.5
Gordon15 | 1 | 1 | visual | FastICA | 2.5
Gordon15 | 1 | 1 | visual | JADE | 2.5
Gordon15 | 1 | 1 | visual | SOBI | 5.5
Gordon15 | 1 | 1 | visual | PCA | 5.5
Snyder15 | 1 | 1 | Mutual Info | AMICA | 1 | doing something is better than doing nothing ;)
Snyder15 | 1 | 1 | Mutual Info | Original EEG | 2
Snyder15 | 2 | 1 | Mutual Info | AMICA | 1
Snyder15 | 2 | 1 | Mutual Info | Original EEG | 2
Snyder15 | 3 | 1 | Mutual Info | AMICA | 1
Snyder15 | 3 | 1 | Mutual Info | Original EEG | 2
Snyder15 | 1 | 1 | Correlation | AMICA | 1
Snyder15 | 1 | 1 | Correlation | Original EEG | 2
Snyder15 | 2 | 1 | Correlation | AMICA | 1
Snyder15 | 2 | 1 | Correlation | Original EEG | 2
Snyder15 | 3 | 1 | Correlation | AMICA | 1
Snyder15 | 3 | 1 | Correlation | Original EEG | 2
Kline15 | - | - | - | -
Raduntz15 | - | - | - | - | - | new method was compared to experts visual estimates
Mariani15 | - | - | - | - | -
Zhang15 | 1 | 1 | Accuracy | WICA | 1
Zhang15 | 1 | 1 | Accuracy | Original EEG | 2
Zhang15 | 2 | 1 | Accuracy | WICA | 1
Zhang15 | 2 | 1 | Accuracy | Original EEG | 2
Zhang15 | 3 | 2 | Accuracy | WICA | 1
Zhang15 | 3 | 2 | Accuracy | Original EEG | 2
Zhang15 | 4 | 2 | Accuracy | WICA | 1
Zhang15 | 4 | 2 | Accuracy | Original EEG | 2
Bulea14 | 1 | 1 | accuracy | Raw | 1.5
Bulea14 | 1 | 1 | accuracy | ASR | 1.5
Chen14 | 1 | 1 | SNR | Original EEG | 6
Chen14 | 1 | 1 | SNR | SCICA | 4.5 | unable to effectively separate muscular artifacts from brain activity
Chen14 | 1 | 1 | SNR | EEMD-ICA | 4.5 | unable to effectively separate muscular artifacts from brain activity
Chen14 | 1 | 1 | SNR | EEMD-CCA | 2.5 | was able to effectively isolate muscle activity into the bottom components
Chen14 | 1 | 1 | SNR | EEMD-IVA | 2.5 | was able to effectively isolate muscle activity into the bottom components
Chen14 | 1 | 1 | SNR | EEMD-JBSS | 1
Chen14 | 2 | 2 | SNR | Original EEG | 6
Chen14 | 2 | 2 | SNR | SCICA | 4.5 | unable to effectively separate muscular artifacts from brain activity
Chen14 | 2 | 2 | SNR | EEMD-ICA | 4.5 | unable to effectively separate muscular artifacts from brain activity
Chen14 | 2 | 2 | SNR | EEMD-CCA | 2.5 | was able to effectively isolate muscle activity into the bottom components
Chen14 | 2 | 2 | SNR | EEMD-IVA | 2.5 | was able to effectively isolate muscle activity into the bottom components
Chen14 | 2 | 2 | SNR | EEMD-JBSS | 1
Yilmaz14 | - | - | - | - | - | was not compared to anything, but best results obtained when eyes closed and jaw dropped conditions were met
Reis14
Boudet14 | 1 | 1 | Percentage of estimated removed artifacts | DAFOP | 1 | Comparison of this paper seems generally weak
Boudet14 | 1 | 1 | Percentage of estimated removed artifacts | BSS-CCA | 2
Boudet14 | 1 | 1 | Percentage of estimated removed artifacts | LP-FILTERING | 3
Li13 | 1 | 1 | Classification Accuraccy | L1-CSP | 1 | only real dataset accuracies from experiment compared, simulated dataset ignored
Li13 | 1 | 1 | Classification Accuraccy | L2-CSP | 3
Li13 | 1 | 1 | Classification Accuraccy | TR-CSP | 2
Lawhern13 | 1 | 1 | visual | DETECT | 1 | ranking used to rank effectiveness of presented method among different datasets (D1-D3), D1 had signifacently less artifact contamination
Lawhern13 | 1 | 1 | visual | DETECT | 2.5
Lawhern13 | 1 | 1 | visual | DETECT | 2.5
Nottage13 | - | - | - | - | - | no comparison needed
Yong12 | 1 | 1 | TPR | Ignore | 10 | for 0.0s dwell time
Yong12 | 1 | 1 | TPR | Reject | 7
Yong12 | 1 | 1 | TPR | SOBI | 5
Yong12 | 1 | 1 | TPR | ERICA | 8
Yong12 | 1 | 1 | TPR | AMUSE | 6
Yong12 | 1 | 1 | TPR | SWT-SURE | 9
Yong12 | 1 | 1 | TPR | ASWTD-Hard | 4
Yong12 | 1 | 1 | TPR | ASWTD-Soft | 2
Yong12 | 1 | 1 | TPR | ASWTD-Garrot | 1
Yong12 | 1 | 1 | TPR | ASWTD-SBSS | 3
Yong12 | 2 | 1 | TPR | Ignore | 10 | for 0.25s dwell time
Yong12 | 2 | 1 | TPR | Reject | 7
Yong12 | 2 | 1 | TPR | SOBI | 5
Yong12 | 2 | 1 | TPR | ERICA | 8
Yong12 | 2 | 1 | TPR | AMUSE | 6
Yong12 | 2 | 1 | TPR | SWT-SURE | 9
Yong12 | 2 | 1 | TPR | ASWTD-Hard | 4
Yong12 | 2 | 1 | TPR | ASWTD-Soft | 2
Yong12 | 2 | 1 | TPR | ASWTD-Garrot | 1
Yong12 | 2 | 1 | TPR | ASWTD-SBSS | 3
Yong12 | 3 | 1 | TPR | Ignore | 7 | for 0.75s dwell time
Yong12 | 3 | 1 | TPR | Reject | 10
Yong12 | 3 | 1 | TPR | SOBI | 5
Yong12 | 3 | 1 | TPR | ERICA | 8
Yong12 | 3 | 1 | TPR | AMUSE | 3.5
Yong12 | 3 | 1 | TPR | SWT-SURE | 9
Yong12 | 3 | 1 | TPR | ASWTD-Hard | 6
Yong12 | 3 | 1 | TPR | ASWTD-Soft | 1
Yong12 | 3 | 1 | TPR | ASWTD-Garrot | 2
Yong12 | 3 | 1 | TPR | ASWTD-SBSS | 3.5
Yong12 | 4 | 1 | MSE | SOBI | 6.5
Yong12 | 4 | 1 | MSE | ERICA | 8
Yong12 | 4 | 1 | MSE | AMUSE | 6.5
Yong12 | 4 | 1 | MSE | SWT-SURE | 5
Yong12 | 4 | 1 | MSE | ASWTD-Hard | 2.5
Yong12 | 4 | 1 | MSE | ASWTD-Soft | 2.5
Yong12 | 4 | 1 | MSE | ASWTD-Garrot | 2.5
Yong12 | 4 | 1 | MSE | ASWTD-SBSS | 2.5
Yong12 | 4 | 1 | Spectral Distortion | SOBI | 7.5
Yong12 | 4 | 1 | Spectral Distortion | ERICA | 5
Yong12 | 4 | 1 | Spectral Distortion | AMUSE | 7.5
Yong12 | 4 | 1 | Spectral Distortion | SWT-SURE | 6
Yong12 | 4 | 1 | Spectral Distortion | ASWTD-Hard | 2
Yong12 | 4 | 1 | Spectral Distortion | ASWTD-Soft | 4
Yong12 | 4 | 1 | Spectral Distortion | ASWTD-Garrot | 2
Yong12 | 4 | 1 | Spectral Distortion | ASWTD-SBSS | 2
Lau12 | - | - | - | -
Lawhern12 | - | - | - | -
Winkler11 | 1 | 1 | Cross-validation error | LPM-Setup1 | 2.5 | Only setup 1 is comparable by Cross-validation error
Winkler11 | 1 | 1 | Cross-validation error | SVM-Setup1 | 1
Winkler11 | 1 | 1 | Cross-validation error | LDA-Setup1 | 2.5
Winkler11 | 1 | 1 | Cross-validation error | TrainMethod-Setup 2 | -
Winkler11 | 1 | 1 | Cross-validation error | CSP-LDA-Setup 3 | -
Shackman10
Gwin10 | 1 | 1 | spectral power | Original EEG | 3
Gwin10 | 1 | 1 | spectral power | channel-based artifact removal | 2
Gwin10 | 1 | 1 | spectral power | IC-based artifact removal | 1 | much more complexity, but only slightly (but significantly) better
McMenamin10 | - | - | - | - | -
Logesparan10
Mosher09 | - | - | - | - | - | No comparison possible, only new method is being shown
McMenamin09 | - | - | - | - | -
Delorme04 | - | - | - | - | - | No comparison is needed for this paper, EEGLAB is being explained
Iriarte03 | 1 | 1 | Visual | JADE-ICA | 1
Iriarte03 | 1 | 1 | Visual | HF_LF_NOTCH | 2
Jung01 | - | - | - | - | -

## Devices
abbreviation | name
EEG | electroencephalogram
MEG | magnetoencephalogram
EMG | electromyogram
ECG | electrocardiogram
IMU | inertial measurement unit
SMU | source measure unit
Reference | Sensors & Electrodes
Most papers | Ag-AgCl (gel-electrode)
Reis 14 | MEMS (dry micro-electromechanical system sensors)
Reis 14 | Photrodes (Srico, elector-optical)
Reis 14 | g.SAHARA (g.tec)
Reis 14 | Cognionics (Cognionics)
Reis 14, Grosselin19 | actiCAP electrodes (Brain Products)
Wang17 | Printed dry concentric Electrodes (own)
Tamburro18, Stone 18 | Multipin Polyurethane electrodes
Zhang20 | 3D printed PWS electrode coated with PEDOT:PSS
Beach21 | IMU with accelerometer and gyroscope on each electrode and using AF to reduce motion artifacts directly on the active electrode
Grosselin19 | g.LADYbird (g.tec)
Grosselin19 | wetRS
Grosselin19 | dryRS
Chu18 | Ag/AgCI passive gel electrodes (Plexon Inc., USA)
Arad18 | eego sports 32 pro by ANT-Neuro
Reference | Caps & Amplifiers
Most papers | ANT Neuro cap & amplifier
Reis 14 | BrainWave (Medi Factory, Netherlands)
Reis 14, Stone18 | WaveGuard (ANT-Neuro, Netherlands)
Reis 14 | Brain Vision, actiCAP (Brain Products)
Reis 14 | EEG headset (Cognionics, USA)
Reis 14 | MOVE system (Brain Products)
Reis 14 | eegosports (ANT-Neuro)
Reis 14 | Cognionics wireless EEG acquisition unit
Reis 14 | Mindo (National Chiao Tung University)
Jung01 | Electrocap (Electro-Cap International Inc, USA)
Gwin10, Symeonidou18, Lawhern13, ... | 248-channel active electrode array (BioSemi ActiveTwo, Netherlands)
Tamburro18 | ECI-Electrogel, Electrocap International Inc.
Tamburro18 | RefaExt (ANT)
Anders20 | 64-channel active wet Ag/AgCl electrodes on an elastic cap (Easycap, Germany)
Anders20 | Live Amp, Brain Products, Germany
Li18 | Neuracle 8-channel wireless EEG (Neuracle, China)
Snyder15, DelPozo-Baños17, Symeonidou18, Lawhern13, ... | 256 channel, BioSemi Active 2 (backpack)
Snyder15, DelPozo-Baños17, Symeonidou18, Lawhern13, ... | BioSemi Active 2
Tost21 | Starstim 20 wireless device (Neuroelectrics, Spain)
Wang17 | OpenBCI board set
Zou16 | Custom hardware with TI ADS1299 that can measure the impedances of the scalp-electrode contacts with helps to detect non-biological artifacts
Delisle-Rodriguez17 | Synamps2 (Neuroscan)
Delisle-Rodriguez17 | BrainNet BNT 36 (EMSA, Brazil)
Rosanne21 | Neurolectrics Enobio EEG System
Grosselin19 | Melomind (myBrain Technologies, Paris, France)
Grosselin19 | g.USBamp (g.tec)
Grosselin19 | Acticap (BrainProducts)
Grosselin19 | GAMMAsys (g.tec)
Chu18 | OmmiPlex Neural Data Acquisition System (Plexon Inc., USA)
McMenamin10 | Geodesic Sensor Net (GSN128, Electrical Geodesics Inc. USA)
Iriarte03 | LaMont amplifiers (LaMont Medical, Madison, WI, U.S.A)
Costa16 | g.GAMMAbox, g.Tec, Austria
Yong12 | L64 Sagura EEG amplifier system
Arad18 | eego sports 32 pro by ANT-Neuro
Gabard-Durnam18 | HydroCel 128-channel Geodesic Sensor Nets (EGI, Eugene, OR)
Blum19 | Measurement on a Smartphone (Sony Xperia Z1), running mBrainTrain App with a EEG cap
Leach20 | EGI Netstation software (Electrical Geodesic Inc.)
Leach20 | EGI NetAmps 300 (Electrical Geodesic Inc.)
Melman16 | FS128 Toolbox, XLTEK acquistion system (Natus Medical, Canada)
Soghoyan21 | NeuroTravel amplifier (EB Neuro, Italy)
Yilmaz14 | MEDCAP, SpesMedica S.r.l, Italy
Oliveira16 | Cognionics Mobile, Cognionics Dry
Weiss17 | "EEG-1200" amplifier (Nihon-Kohden, Irvine, CA)
Jas17 | Neuromag VectorView system
Jas17 | BCI2000 EEG cap
Jas17 | Easycap EEG
Reference | Other new device setups
Snyder15 | novel 3 layer system: [1] Silicon swim cap; [2] simulated scalp, [3] EEG system
Nordin20, Nordin19 | novel 2 layer EEG: [1] 128-scalp EEG electrodes with BioSemi ActiveTwo sensors [2] custom conductive fabric cap (Eeonyx, Fig. 1B) which approximately matched the resistivity of human skin Dual layer EEG simultaneously recorded human electrocortical signals and isolated motion artifacts using pairs of mechanically coupled and electrically independent electrodes and a custom conductive fabric cap
Arad18 | Bi-lateral force plates (R-Mill, ForceLink, NL)
Yilmaz14 | SMU measurement with intramuscular silver fine-wire electrodes coated with Teflon

## Motion Artefact Removal
abbreviation | name | variants | details or usage in pipeline
Filters
Filtering | Filter (HP, LP, BP, BR) | Butterworth, Bessel, Tschebyschef, Finite Impulse Repose (FIR) | BP 0.1-40 (FIR): means a band-pass with 0.1Hz and 40 Hz cuttoff-frequencies using the filter FIR. The transition band (signal output is between 10% and 90%) is not specified
MA | Moving Average | Moving Mean (standard), Moving Median, Weighted Moving Average
Other | Robust Detrending | trial-masked robust detrending, robust detrending | Original from Driel et. al (2021, [77]), was dismissed because no motion focus but shows to have smaller impact on original signal compared to traditional HP filtering for removing baseline drifts. Could be interesting for us, but paper could not be added to PASS folder as it does not meet filter ruleset requirements.
Aggregation
Epoching | Epoching into time intervals / windows | Sliding Windows
Feature Extraction | Feature extraction | Used features: Range, Kurtosis, Standard deviation, Autocorrelation, Focal Components, Correlation between channels, CCA-components, fractal dimension, TODO: find all used features in the pipeline | Extract features from all data or time windows
BSS (blind source separation), Decomposition and Dimensionality Reduction
FT | Fourier transformation | FFT (Fast FT), STFFT (short-time) | In sliding windows FFT + Cancel bad coefficients + Inverse FFT (STFFT?)
Wavelet | wavelet package decomposition | DWT (discrete wavelet transformation), SWT (stationary wavelet transformation), Wavelet Denoising, xDAWN, Gabor functions (e.g. as Gabor wavelets of for Gabor Fitting) | Different Wavelet families are: Daubechies, Symlet, Coiflet and Fejer-Korovkin. SWT seems to be more robust, as no interpolations are needed
Welch | Welch power spectral density
LSP | Lomb-Scargle periodogram |  | tampering in overlapping sliding windows and then FFT
PCC | Pearson Correlation Coefficient | CCC (concordence correlation coefficient)
CCA | canonical-correlation analysis | GECCA (Gaussian Elimination CCA) , MCCA (multiset CCA)
ICA | indipendent component analysis | wICA (wavelet-thresholding ICA, Geetha, 2011, ICA-W is maybe the same), SASICA, Fast-ICA, AMICA (Adaptive Mixture Independent Component Analysis), Merged AMICA, Infomax ICA, JADE, EAWICA (First wavelet, then ICA on wavelets, Mammone 2014), ERICA (Equivalent Robust ICA, based on cumulants, Yong12), CUDAICA (Weiss17) | after wICA ICA is applied, because it improves the final ICA decomposition (Gabard-Durnam, 2018). ICA is computationally expensive and not good for online studies. AMICA and InfomaxICA seems to be the best suited ICA algorithm says Delorme.
NMF
PCA | Principal Component Analysis | t-PCA, ASR (artifact subspace reconstruction), rASR (Riemann geometry ASR) | PCA is a dimensionality reduction method that calculates (orthogonal) eigenvectors and weights those components based on their eigenvectors. ASR is the most common and widespread online method of removing transient, high-amplitude-related artifacts from different sources like eye blinks, muscle bursts, and movement [26] while recovering essential EEG background activities that lie in the subspace spanned.
t-SNE | t-distributed stochastic neighbor embedding
EMD | empirical mode decomposition | EEMD (ensemble empirical mode decomposition), sEEMD (signal serialization based), Noise-assisted Multivariate Empirical Mode Decomposition (NA-MEMD) | Iteratively calculates the mean of the minimal and maximum envelope of a signal as IMF-components (intrinsic mode functions) and subtracts them from the signal to get a new one [Roy17].
SSD | spatio-spectral decomposition |  | SSD aims to extract oscillations in a frequency band of interst at maximal singal-to-noice ration.
JBSS | joint blind source separation | JBSS | separating a set of source signals from mixed signals without the aid of information about the source signals or the mixing process
IVA | independant vector analysis | IVA | In independent vector analysis (IVA), the goal is mixture and source identification (BSS) for a collection of disjoint but coupled data sets. Within the same data set, the source signals are assumed to be statistically independent, while between data sets, source signals may be correlated. IVA can thus be seen as a multi-set generalization of independent component analysis (ICA)
CSP | common spatial patterns | Filter-Bank CSP | CSP is a method used to decompose multivariate signals into additive components such that it maximazis the variance
SSA | singular spectrum analysis |  | Decomposes a single-channel signal into eigenvalues and eigenvectors and reconstructs those with a autocorrelation below a given threshold
Other | Delisle-Rodriguez17: 4: Local Average Reference 4: Weighted Average Reference 4: Riemann Kernel Yong12: SOBI (Second order Blind Identification) AMUSE (Algorithm for Multiple Unknown Source Extraction) Lawhern12: Auto-regressive model Lau12: Weighted Phase Lag Index. Assumes that a 90-degree out-of-phase signal arises only from highly complex coupled harmonic oscillators like the brain and not from artifact sources. Is based on the fraction of the Hilbert transform and the input itself Jas17: SNS - Sensor Noise Supression: reduces the dimensionality of multichannel data
Artifact Detection or Classification of Motion Artifacts
Linear Regression | Linear regression | robust Linear Regression, Polynomial Regression | Used to calculate residuals of individual epochs
DA | discriminative analysis | LDA (linear discriminant analysis), LFDA (Local Fisher's DA) | Discriminate Classes based on their inter-class and intra-class variablity
SVM | Support Vector Machine | OCSVM (one class)
sph. spline | spatial spherical spline | spatial spherical spline | A reconstruction method used for the interpolation of bad channels based on spatial information of neighbored channels. Perrin et al., 1989
GMM | gaussian mixture model |  | used for unsupervised learning. It is a probabilistic model that assumes that all data points are generated from a mixture of a finite number of Gaussian distributions with unknown parameters
Saini-VMD (no name given) | [1] VMD; [2] Zero crossing count threshhold; [3] Rejection/Retention | newly proposed, so far no variants found, uses VMD approach analyzing two modes of the signal | A single channel EEG method used to identify artifact components without any other input needed
CSP | Common spatial patterns | L1-CSP, L2-CSP, TR-CSP | maximize variance of one class, minimize the variance of all others
Other | RELICA, BeamICA, ICLabel are plugins of EEGlab that can classify the types of ICs like brain activity, muscle, heart, eye, noise; hierarchical clustering can be use to detect artifactual epochs from extracted features From Jas17: + Autoreject + RANSAC
Artifact Correction
Autoencoder | Autoencoder | VAE | Encoder-Decoder Architecture for reconstructing the artifact-contaminated signal
GAN | Generative Adversarial Network | GAAL, MGAAL
Other | multichannel EEG thresholding by similarity | METS | see Saavedra19; #171
| SuBAR | SuBAR | Surrogate Based Artifact Removal
| generalized sidelobe canceller | GSC | see Mosher09; #640
Special Methods for EEG
MARA |  | iMARA (infant, specialized for infant EEG, Haresign et al., 2021) | For the automatic rejection of components associated with artifacts. Often coupled with ICA. Winkler et al., 2011
ACEMIC | Automatic Classification of Electromyogenic ICA Components |  | As MARA, it automatically detects and removes IC originating from muscle artifacts (Gabsteiger et et. 2013)
DSS | Denoising Source Separation |  | A method to clean ERP. Used after other filterings before the data averaging. De Cheveigne and Parra (2014), de Cheveigne and Simon (2008)
DFA | Detrended Fluctuation Analysis
AF | Adaptive Filtering
FASTER | fully automated statistical thresholding for EEG artifact rejection |  | Not working channels are identified using FASTER. Nolan et al., 2010
ADJUST | Automatic EEG artifact Detection based on the Joint Use of Spacial and Temporal features | adjusted-ADJUST (, specialized for developmental data, Leach et al., 2021) | Artifact Removal algorithm. Mognon et al., 2011
ERASE | EMG Removal by Adding Sources of EMG
STAR | sparse time artifact removal |  | Removes artifacts that are sparse in space and time
ARNG | amplitude range knee point in gait cycle |  | To find channels with a high amplitude range within epochs
HAPPE | Harvard Automated Processing Pipeline for Electroencephalography |  | pipeline suitable for low density EEG channels and limited data samples. It relies on wICA and multiple artifact rejection algorithm (MARA) to detect artifactual components for rejection.
LCF | Localized component filtering | LCF | can be applied to all existing BSS systems as a post-processing-step as it uses BSS components as inputs and returns "cleaned" BSS components. Internally uses Feature Extraction, an Integrator, a Decision Logic and a Mixer component. For more information see #333 (DelPozo-Baños et. al)
Other | semblance-based ERP window | SEWS-1, SEWS-2, The Fingerprint Method (described in e.g. Stone18) | can most likely only be used upon ERP classification problem
clean_rawdata (EEGLAB function)
LORETA (low-resolution electromagnetic tomography, McMenamin)

## Classification Models
abbreviation | name | variants | assumptions | details or usage in pipeline
DA | discriminative analysis | LDA (linear DA), QDA (quadratic DA), PSD comparison with threshhold
KNN | k-nearest neighbors | KNN
KMC | k-means clustering | KMC
ED | Euclidean distance | EUC
NB | Naive Bayes
DT | Decision Tree
RF | Random Forest | IF (isolation forest)
SVM | Support Vector Machine | OCSVM (one class)
NN | Neural Network | RBF (Radial Basis Function), MLP (multi-layered perceptron), CNN (convolutional NN), DBN (deep belief network), restricted Boltzmann machine (RBM), ... DCNN - Deep Learning Convolutional Neuronal Network (see #82, Namias et. al, 2021 for further reference)
GMM | gaussian mixture model |  |  | It is a probabilistic model that assumes that all data points are generated from a mixture of a finite number of Gaussian distributions with unknown parameters. In most cases a very low number of components per subject and class is sufficient (1).
MEC | Minimum energy combination | can be used for classification in EEG |  | MEC combines the data from the least noisy channels. The amount of noise in the data recorded by each channel is calculated by removing the target signals and assessing the rest. The combination is executed for each target frequency, resulting in comparable SNR measures for all targets. After normalization, the result of the classifier is given in percentages for each investigated frequency.
Others:
XGB | extreme gradient boosting
AdaBoost
LPM | Linear Programming Machine

## Software
Name (these frameworks are used within the pipeline) | Frameworks based on the language
Matlab | EEGLAB, ERPLAB, DETECT, ... | MoBILAB, DataSuite
Python | MNE toolkit, Neuro Digital Signal Processing Toolbox (NeuroDSP), PyEEG, scikit-learn
Other services |  | Cartool (swiss), EMSE, LORETA-KEY (swiss), SPM5 (UK)
Other (not used within the pipeline)
SPSS | e.g. IBM SPSS statistics 21 (IBM) | for statistical analysis
R |  | for statistical analysis
Statistica
Neuroscan | Neuroscan 4.3 software (e.g. Nottage13)
Eevoke | ANT Neuro | for creating EEG scenarios
BEAPP | Batch EEG Automated Processing Platform | Include different EEG channel layouts within a single pipeline and run the HAPPE algorithm
BESA Dipole Simulator | MEGIS Software GmbH, Gräfeling, Germany | For the simulation of dipoles in the brain and expected electrode potentials
Spike2 | Cambridge Electronic Design | Used for analysis of EEG records in Yilmaz14
SystemPlus | Micromed | Used as recording software in Yilmaz14
BrainRecorder Software | BrainProducts GmbH | used for BrainProducts systems for recording
E-Prime application suite | Psychology Software Tools | To create multiple tasks for BCI experimental design
VS.NET | Multi-Attribute Task Battery (MATB-II) | used for triggering mental workload and simultaneous workload assesment
Harmonie | Harmonie software (Stellate, Montreal, Quebec, Canada) | https://www.medicregister.com/Stellate_Systems_Inc/Supplier/sid1448.htm
TracerDAQ | TracerDAQ software (National Instruments, Austin, TX, USA) | Data acquisition (DAQ) software provides four virtual instruments to acquire, analyze, and display data, and generate signals.
Persyst v12 | Persyst GmbH | artifact reduction software tool (also see #Weiss17): https://www.persyst.com/technology/artifact-reduction/
Motion Analysis |  | for MOCAP
Visual-3D | C-Motion | Motion Capture System
Vicon Nexus | Oxford, UK | Motion Capture System
FaceLAB eye tracking system | Facelab | Eye tracking EOG system
http://www.neurology-kuleuven.be/ | Open Software technology provided by De Clercq et al | Used by Boudet14, but doesnt come with an english version?
Brainstorm | for details see: Tadel et al., 2011 | rejection of bad data segments (also see #Jas17)
FieldTrip | for details see: Oostenveld et al., 2011 | rejection of bad data segments (also see #Jas17)
MNE | for details see: Gramfort et al., 2013 | rejection of bad data segments (also see #Jas17)
SPM | for details see: Litvak et al., 2011 | rejection of bad data segments (also see #Jas17)
Open Dataset | (not complete)
Temple University | Temple University EEG Corpus - Downloads (piconepress.com)
SEED | SEED Dataset (sjtu.edu.cn)
PhysioNet | PhysioNet Index
ALICE crowdsource label platform | ALICE.adase.org/ | Soghoyan21
ICLabel crowdsource platform | SCCN: Independent Component Labeling (ucsd.edu)

## Consensus
Referenences | Statement | Consequences | Category
Pretty much all | motion artifacts prohibit the usage of mental signals via EEG in the real world | The research for artifact removal and correction has to accelerate | general
Most authors | Muscle artifacts originating from whole body movement are more complex to handle than other EEG artifact types due to them impacting almost all EEG local subbands with their broad spectrum, high amplitude and high sensitivity. Artifact amplitude typically being higher than signal amplitude is an additional problem. | Multiple different approaches have to be compared or even combined to solve the problem, most likely no single method alone will do the trick | general
Snyder15, Gwin10, Symeonidou18 | There is a lack of methods that separate non-neuronal motion artifacts from neuronal activity (true brain components) | Alternative methods are needed for the identification and separation of movement artifact in mobile EEG signals, especially methods that can be performed in real time. Separating true brain signals from motion artifact could clear the way for EEG brain computer interfaces for assistance during mobile activities, such as walking. Also see: artificial brain setups used in the stated studies | general
Jung01, Tamburro18, Iriarte03 | ICA has very strong assumptions: There are 4 assumptions underlying ICA, which can not assumed in practise: 1) Signal conduction times are equal and the summation linear, i.e. no significant wave time of potential. 2) Spatial projections of components are fixed in time and conditions. 3) Temporal independence of the components, this is not met for motion imaginary tasks that occur both myogenic and mentally. 4) Statistical distributions of the component activation are not Gaussian, therefore ICA requires a longer time windows. Generally, ICA is still very effective and stable for EEG and fMRI. | ICA might not be suited to filter very complex artefacts that reserve a high number of IC. | ICA-limitation
Oliveira17, Jung01, Saba-Sadiya21, Li18, Mur19, Tamburro18, Chen14, Delisle-Rodriguez17, Beach21, Sebek18, Iriarte03 | ICA is limited in the number of independent components. | Motion artifacts are very uncorrelated to each other and therefore occupy a very high number of ICs resulting from the ICA method, thus reducing the IC containing useful mental signals. ICA require a high number of electrodes. "it is practically impossible to obtain a mixing matrix from a training database that contains all possible physiological and non-physiological artifacts, as well as several internal states of the user, due to changes in cognitive-motor-affective behaviors, medication and health status" Delisle-Rodriguez17 cites - Kilicarslan, A.; Grossman, R.G.; Contreras-Vidal, J.L. A robust adaptive denoising framework for real-time artifact removal in scalp EEG measurements. J. Neural Eng. 2016, 13, 026013. | ICA-limitation
Flo22, Saini20, Lawhern13 | Widespread use of portable, high-throughput sensors has dramatically increased the challenges of analyzing time series data, making automated methods for time series analysis not only beneficial but necessary for practical data analysis | More automated pipelines are needed for future work, manual or semi-manual solutions become less future-proof | general-BCI
Nottage13, Arad18 | Most of the previously widely reported induced broad-band gamma peak at 200–300 ms after a visual stimulus does not originate in the brain, but stems from sources of artifacts outside of the brain. The results show that the P300 peak has a high inter- and intra-subject variance, indicating that this is not a good measure for the oddball experiment-Lau12 | The gamma band consists of mainly artifact activity and therefore could be used for artifact detection, artifact correction still needed in alpha/beta band | general-BCI
Val-Calvo19, Nahmias21, Saba-Sadiya21 | There is a lack of long-term deep learning strategies applied to artifact detection/correction/mitigation/removal | Most work focuses on ICA optimization approaches, though optimization of learning algorithms (supervised to unsupervised) might be able to achieve better individual results in the long run / long usage. | ML

## Research Gap
Required | Details | Authors | Category
There is a lack of standardized preprocessing (ICA, etc.) steps that are validated and include a basic filtering of noise signals. | Hardly any theoretical or practical approaches were conducted that examines the affect of filtering methods on the latency and signal form of the mental EEG signals. | Karpiel21: weak results but still done, Anders20 | general
More research in the automated correction of channels and time interval that suffer from motion artifacts, instead of removal which leads to a loss of information during artifact-affected intervals | Only few publication provided a pipeline that corrects motion artifacts | Anders20, Yong12, Zhang15 | general
A very harsh averaging is applied across gaits, epochs, trials, and even subjects | This is often performed in order to demonstrate "typical brain patterns". However, the great inter-subject variability shows that this explanation could be missleading and this approaches doesn't help to build a BMI | Kline15, Nathan16 | general
We found very rare mass benchmark testing papers were a new or existing algorithm or pipeline was evaluated over many different databases, mental strategies, participants or system setups | A rare exception for this was "Autoreject", presented in Jas17, whisch was evaluated against 4 different databases with oover 200 subjects data | Authors, (ref: Jas17) | general
There is the need for pipeline working online in an artifact-rich EEG. Many pipelines have very strong assumptions on a scarcity of artifacts. |  | Mur19: artifact-free time before and after each artifact and no bad electrode, Cheveigne16: Does not work if an artifact affects multiple electrodes, Jung01: Clearifies assumptions all ICA-methods found are based on. | general
Motion artifacts have traditionally drastically limited the conditions in which scientists have studied human brain dynamics. Experiments are created to avoid these factors rather than progressing on solving these issues | By using combined hardware and signal processing solutions for motion artifact removal, Nordin19 found it is possible to identify human brain activity while humans stepped over obstacles while walking and running (online scenario working!). According to PubMed (state 2019), there were over 2,800 studies on human EEG published in 2017, yet less than 1% were on mobile subjects. Definitely needs more research tackling the issues for real life EEG scenarios! | Nordin19 | general
Single-Channel dry electrode subject-driven classification performance with LDA, Linear SVM, Euclidean or Weighted kNN needs long lasting individual studies with longitudinal recordings (weeks or months) to reach its full performance potential | The presented algorithm in Grosselin19 is not subject-driven but the authors noted there could be great potential for the compared subject-driven methods (LDA, Linear SVM, Euclidean or Weighted kNN) for optimization of classifiers by using individual longitudinal recordings (weeks or months) | Grosselin19 | general
Existing ICA-based removal strategies depend on explicit recordings of an individual’s artifacts or have not been shown to reliably identify muscle artifacts. | Claims that a universal and efficient classifier of ICA components for the subject independent removal of artifacts from EEG is still missing | Winkler11 | general-artifact-removal | has changed
In "Winkler11" succesful removal of few artifacts and correctly detected outliers did not lead to better motor imagery classification accuracy (CSP used). It seems to remain unclear why... maybe because the filtering remove characteristic brain signals essential for the downstream task | Problem addresses the potential performance increase by careful artifact removal for later classification. In our literature list there are few or no papers that investigate the maximum potential impact of muscle artifact removal on later classification outcome, when using common spatial patterns -> suggests to investigate wether correctly detected artifacts can even always lead to better motor imagery classification accuracy | Winkler11 | Ground-truth-problem
The performance evaluation of artefact removal methods is problematic due to the lack of standard quantitative metrics in the literature to measure both the amount of artefact removal and the distortion [1,26,27] | Most of the methods are often tested on real EEG, and their performances are evaluated in terms of some qualitative plots analysed by visual inspection. Investigate who classifies based on the corrected EEG Lack of community agreed-upon benchmarks for comparing the performance and usability of EEG-systems. Our results suggest that the use of SNR alone as a metric for benchmarking may generate questionable results, but the combination of SNR and CV can establish a robust method.Oliveira16.-Naja keine Ahnung was da wirklich validiert wird. | Mur19, Oliveira16 | Ground-truth-problem
There is few literature that corrects real artifacts, as the ground true can't be measured directly. | "Another common limitation is the prevalent use of simulated artefacts for the validation of the classification system; however, simulated artefacts cannot represent the entire variety of real artefacts possibly affecting brain activity recordings (e.g., Barbati et al., 2004; Halder et al., 2007; Nolan, Whelan & Reilly, 2010; Hou et al., 2016)" | Tamburro18, Yong12 | Ground-truth-problem
Single motor unit studies (focusing on one individual muscle or muscle group) with an attempted ground truth measurment have a very high value as they add knowledge of individual muscle contribution to the EEG system. There should be more studies targeting different muscles or muslce groups, but they seem to be rare. | In Yilmaz14 we can find information on how the temporalis muscle impacts the EEG. Understanding EEG contamination on the level of individual muscle contribution helps identify new ways to detect and mitigate their effects. We have found very few papers in our research that try to separate EEG muscle artifact contamination into contribution of individual muscles / muscle groups. In almost all the found work, all muscle contamination is seen as a summed up artifact that gets detected and removed as a whole at the same time. This approach might be the wrong way to do it though. | Yilmaz14 | Ground-truth-problem
Adding measured EMG artifact channels to EEG channels improves ICA, as it learns and detects EMG contamination within EEG forcing EMG artifacts into a minimal number of ICs. Different methods seem to use different ways of mixturing these EMG signals into EEG channels | Comparision of EMG-added removal methods to understand optimal mixture of signals from EMG and EEG | Li21 | Ground-truth-problem
Beneficial effects of placing additional EMG electrodes over the face and neck muscles should be tested with different algortihms (in different combinations of EMG channels) of this pipeline to see which methods benefits most | Only EEMD-CCA was tested on improval | Mucarquer20 | Ground-truth-problem
There is a lack of standard metrics that evaluate the artifact removal on a downstream task. The methods are often too aggressive and also filter mental signals to a severe or unknown degree. | Nearly all investigated papers evaluate their pipelines by comparing corrected EEG signals. This approach, however, does not take into account whether the brain signals for interest is retained in the corrected signal, or the signal is even removed among myogenic artifacts. In order to evaluate the retainment of the brain signals, an additional downstream task with an appropriate mental strategy has to be defined in the testing setup. It has to be taken into account, that this approach can only evaluate the method's fit of purpose, i.e. the ability to correct artifactual noise while retaining brain signals for a specific task and is therefore constrained to an defined use-case. Delisle-Rodriguez17 compared the WAR with and without Adaptive filtering, and with adaptive filtering there were much less attenuation, but the performance was still worse. | All but those papers that have a real classification model | method-comparision
There are hardly any comparisions between well established ICA variants and competitive algorithms. | Only wICA was compared with standard ASR so far. |  | method-comparision
The Fingerprint Method needs comparison studies with other existing artifact removal methods (especially interesting in the sports domain). | Stone18 was limited to addressing only the performance improvement of Fingerprint Method. It remains an open question how the Optimized Fingerprint Method directly compares to the method of Gabsteiger et al or other methods of automatic artifact removal in sports science | Stone18 | method-comparision
There are hardly papers that compare already existing methods to each other. Even worse, most papers compare their proposed method only to a default pipeline without applying any method to remove motion artifacts. It is highly suggested, that any newly proposed method should be compared quantitatively to already existing and successful pipeline in order to increase the comparability and suited use-cases of the pipelines. |  | All but: Saba-Sadiya21, Flo22, Frohlich17, Karpiel21, Yong12 | method-comparision
Use the principle of a socalled "antagonists advocate" for implementing existing approaches. Students could be told to implement competitive approaches using standard methods. This will increase the robustness of the proposed algorithm in the comparison. | The proposed pipeline becomes much more effort during the implementation and therefore has an advantage compaired to a classical pipeline that is used without e.g. appropriate preprocessing. | A lot of paper only compare against a simple implementation | method-comparision
A parallelization of the pipeline into multiple streams with different preprocessing and therfore a different variant of the EEG data could increase the robustness and decrease the variance of the results. | It is noteworthy that especially these parallel streams within a pipeline could not be adequately represented in the used pipeline description. In a future work, a more sophistiacated approach to represent a pipeline as graph with methods as nodes and the directed edges as data views forwarded. | Tost21, Flo22, | method-comparision
The muscles in the neck compensate head movements during full body movements and therefore EEG is time-shifted to the acceleration | This phase shift has to be taken into account for models that try to incorporate the acceleration of the head. This could be done via CNNs with the acc as additional input or with the full wavelet coherence of the acceleration. Standard Decomposition Methods like ICA, ASR are not suited to model this phase delay | Ideas of the authors based on an issue described in Nathan16, Kline15 | method-comparision
An adaptation of the effective ICA method such that the number of IC can be larger than the number of channels. That would assume the mixing matrix W to be non-quadratic and the inverse could be solved using the Moore-Penrose Pseudoinverse of W. The other option could be that ICA is applied iteratively as long as mainly artifactual components can be removed. | Consensus that ICA is limited in the number of IC. Anders20 indicates that tasks that involve more and more heterogeneious movements reserve more ICs of movement artifacts, thus reducing the ICs capable of filtering solely brain signals. Artefacts reservate ICs. If the artifacts are frequent, heterogenious and high in variance (walking, full body movement), a very high number of ICs is reservated for them and the only a small fraction of ICs can represent brain signals. In Anders20, in the walking task, in averarge only 25 ICs reflect brain signals. "As mentioned above, the benefit of removing artifacts by ICA is limited by the fact that most artifacts in newborn EEG data are non-stereotyped, therefore not easily captured by ICA. Nonetheless, ICA (and in particular Adjusted-ADJUST (Leach et al., 2020), the IC classifier developed for infant data and included in MADE) may bebeneficial as a further processing step after NEAR because it could correct the data from residual stereotyped artifacts without any further data rejection." Kumaravel22 |  | ICA-limitation
More combinations of methods for BSS are needed that can maximize the muscle activity of groups of muscles into well defined components. The ICA permutation problem in combination with groups of muscles involved in one activity produces worse results on existing BSS, which remains a gap in existing methods (Chen14: EEMD-JBSS?) | When muscular artifacts involve the simultaneous movement of a group of muscles (without stereotyped topography), multiple ICA-based BSS methods are unable to find components without muscle activity. All components instead show both: brain activity and muscle activity. | Chen14 | ICA-limitation
Research is needed that solves the maximum IC limitation problem, restricting the max. obtainable ICs from ICA to channel size. Moore-Penrose Pseudoinverse should be tried out in different combinations with different ICA strategies with direct performance comparison | Moore-Penrose Pseudoinverse can be used to solve the ICA component limitation problem: Matrix division requires quadratic matrices on mixing process (matrix dimensions being channels times ICs). Matrix division means creating the inverse of the matrix and perform a multiplication. Creating an inverse of a matrix is only possible for quadratic matrices. Therefore possible maximum ICs obtained by ICA methods is limited to the number of maximum channels. Channel size restricts max. components. | Sebek18 | ICA-limitation
More research in automatic removal of artifactual components or the correction of motion artifacts using methods without strong assumptions, such as the presence of artifacts is constrained in time or channels | Despite a number of improvements, most pipelines, i.e. the adaption of one pipeline to a new subject, is still semi-automated |  | general-BCI
There is yet no research that investigated alpha and beta power increase/decrease on motion tasks other than treadmill walking; requires more research for collaborative tasks, sport tasks and other domains | For treadmill walking: alpha/beta power increased in contralateral limb single support and push off, and decreased during swing at each gait speed. At faster walking speeds: spectral power fluctuations had limited duration and bandwidth, along with reduced alpha and beta power across the gait cycle, after muscle artifact removal. | Nordin20 | general-BCI
Research is needed looking at somatosensory cortex and motor cortex at the same time analysing spectral power behaviour for tasks that involve greater amounts of sensory feedback built into motor execution | Reduced sensorimotor spectral power could be indicative of greater cortical resources attuned to sensory feedback at faster locomotion speeds | Nordin20 | general-BCI
More electrode comparison studies with different hardware cap setups with and without cable involving real participants in real application domains is required | In Symeonidou18 only a robotic simulation of motion was used in order to focus entirely on the impact of motion artifacts not being disturbed by true brain components. In the future we will want to take these findings back into real measurements with real subjects | Symeonidou18 | general-BCI
Additional studies required comparing wireless systems examining signal quality during head motion in order to determine if wireless systems offer advantages over wired systems in terms of signal quality (wireless VS cable). | In the future, wireless EEG systems might be able to substitute wired systems for locomotion experiments. Paper shows that there are still considerations concerning signal quality, user comfort, and electrode density that need to be addressed in future studies. | Symeonidou18 | general-BCI
For succesful practical implementation of online motor-imagery BCI systems, more studies are needed to solve the decoding performance problem from incomplete EEG signals, rather than fully rejecting heavily contaminated segments (which is a problem for long-term learning strategies) | Consecutive and smooth recognition of SMRs-based BCI systems is needed for the online and long-term application. This requires that the BCI system can continuously decode brain signals without any interruption. If entire EEG segments are discarded due to extreme artifacts or data loss, the BCI system cannot obtain the decoding results during the corresponding time slice. Hence, it is very important to decode incomplete motor imagery EEG for SMRs-based BCI systems in the condition of extreme artifacts and data loss. | Chu18 | general-BCI / ML
Progress in the field of deep learning provides methods from which EEG pipelines could profit. A higher degree of multidisciplinarity, inparticular Machine Learning, could enhance the development of pipelines specialized to correct motion artifacts in EEG signals, similar as active noice cancelling systems became successful. | For instance self-supervised contrastive learning could be used for the pretraining of the network, such that it performs better in the "real" downstream task of e.g. color classification. As the field of EEG analysis is very complex, conducting a machine learning expert for a short time might not lead to the expected results. It is rather suggested to integrate a machine learning expert into the team of EEG analytics such that she or he has sufficient time to really understand the challenges in brain signal analysis. As even experts in other fields struggle in validating what a machine learning expert should provide, the authors would like to give following practical tips: They should have learnt machine learning both in theory (university) and in technological practise (additional hands-on courses for the usage of frameworks) and shall provide practical experiences to cope with time-series data. The papers following papers are the only ones that demonstrate a well-suited usage of state-of-the-art machine learning methods, i.e. methods younger than 10 years. | All but: Saba-Sadiya21 (Auto-Encoder). Some note it as further work suggested: Val-Calvo19, Soghoyan21 Existing Deep Learning CNN for motion artifact correction presented by Nahmias21 (#82) | ML
Usage of evolutionary strategies or genetic algortihms for optimization methods to opt parameters or algorithm parts / pipeline parts from presented list | We have found no evolutionary strategies or genetic algortihms used to optimize the piplines. Pipelines are always static implementations | Mathias Schmoigl-Tonis and Christoph Schranz (idea derived from dismissed paper #684, Gupta07) | ML
Classifiers that are based on features of data decomposition (temporal, spectral, spatial) show a poor generalizability. | There is a need for validation of model transferability to different recording equipment, mental strategies, sessions, subjects, sensor layouts. Maybe it is more successful to use scale-invariant features like parameters from auto-regressive models (Lawhern12) | Tamburro18: refers to Frolich15 | ML
Incorporate participant metadata for the labelling of IC-components. Use more crowdsourcing of artifact labels to validate such models | There exist multiple methods to automatically label the components that origin from ICA decompositions, such as ADJUST, adjusted-ADJUST, RELICA, IClabel, FASTER, MARA, SASICA and BeamICA. To our knowledge, each of them ignores existing and valuable person-metadata such as age or sex, and setup-metadata as EEG type, BCI-paradigm as additional input into their classificator, although we know that the brain signals vary a lot in e.g. children or might be specific to some other context. Modify the ALICE-platform in order to classify artifacts based on the raw EEG-time-series instead of already extracted features. Moreover, publishing also the probabilites for the labels could improve the model training especially for neural networks. Soghoyan21 Better interactivity to discuss epochs difficult to label or very rare artifacts. Use the platform to improve the consistency of the understanding of artifacts in novices and also trained labellers. | cs derived from Leach20 about the adjusted-ADJUST algorithm | ML